## Cover
[6 empty] | DATA SHEET | [5 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  |  | : M0556 |  | MASS FLOWMETER | [5 empty] | DS-IN-F-005
| PROJECT NAME |  |  | : IPA 50KTA Exp. | [7 empty] | REV. NO.
| LOCATION |  |  | : Yeosu, Korea | [7 empty] | A
IPA
| CLIENT |  | LG Chem Ltd.
| PROJECT |  | IPA 50KTA Expansion
| LOCATION |  | Yeosu, Korea
UNIT
DATA SHEET
FOR
MASS FLOWMETER
SERVEONE |  |  |  |  | SERVEONE CONFIDENTIAL INFORMATION: The drawings, designs, ideas, specifications and information contained in this document constitute
confidential information belonging to SERVEONE. Neither this document nor its contents may be used for any purpose or disclosed in any manner
CONSTRUCTION BUSINESS UNIT |  |  |  |  | without the prior written consent of SERVEONE. Reproduction of this document without the prior written consent of SERVEONE or LG Chem.,Ltd.
is prohibited. Any authorized reproduction of this document, or any part thereof, including vendors' shop drawings, shall incorporate this notice."
A |  | MAY.24.'18 |  |  | FOR QUOTAION |  | J.H Jung |  | C.M Lee | K.Y Noh |  |  | H.D Yoon
Rev. |  | Date |  |  | Description |  | Prepared |  | Checked | Approved |  |  | Authorized
JOB NO. |  |  |  |  | DEPARTMENT |  |  | DOCUMENT NO. |  |  | ISSUE DATE
M0556 |  |  |  |  | E&I (INSTRUMENT) |  |  | DS-IN-F-005 |  |  | 2018.05.24

## Mass Flowmeter
[6 empty] | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [11 empty] | CLIENT
[19 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem | [15 empty] | LG Chem
[19 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO. | [15 empty] | DOC. NO.
| JOB NO. |  | : M0556 |  |  | MASS FLOWMETER | [12 empty] | DS-IN-F-005 |  |  |  | MASS FLOWMETER | [12 empty] | DS-IN-F-005 |  |  |  | MASS FLOWMETER | [12 empty] | DS-IN-F-005 |  |  |  | MASS FLOWMETER | [12 empty] | DS-IN-F-005 |  |  |  | MASS FLOWMETER | [11 empty] | DS-IN-F-005
| PROJECT NAME |  | : IPA 50KTA Exp. | [15 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO. | [15 empty] | REV. NO.
| LOCATION |  | : Yeosu, Korea | [15 empty] | A | [16 empty] | A | [16 empty] | A | [16 empty] | B | [15 empty] | A
1 | GENERAL | Tag No. |  |  |  | FT-1111 |  |  |  |  | FT-1129 |  |  |  | FT-1135 |  |  |  | FT-1138 |  |  |  | FT-1139 |  |  |  |  | FT-1319 |  |  |  | FT-1332 |  |  |  | FT-1333 |  |  |  | FT-1336 |  |  |  |  | FT-1350 |  |  |  | FT-1355 |  |  |  | FT-1375 |  |  |  | FT-1401 |  |  |  |  | FT-1402 |  |  |  | FT-1403 |  |  |  | FT-1405 |  |  |  | FT-1435 |  |  |  |  | FT-1839 | [9 empty] | 18
2 |  | P&ID No. |  |  |  | D-1-IPA-1001 |  |  |  |  | D-1-IPA-1001 |  |  |  | D-1-IPA-1001 |  |  |  | D-1-IPA-1001 |  |  |  | D-1-IPA-1001 |  |  |  |  | D-1-IPA-3001 |  |  |  | D-1-IPA-3002 |  |  |  | D-1-IPA-3002 |  |  |  | D-1-IPA-3003 |  |  |  |  | D-1-IPA-3004 |  |  |  | D-1-IPA-3004 |  |  |  | D-1-IPA-3007 |  |  |  | D-1-IPA-4001 |  |  |  |  | D-1-IPA-4004 |  |  |  | D-1-IPA-4004 |  |  |  | D-1-IPA-4004 |  |  |  | D-1-IPA-4003 |  |  |  |  | D-1-IPA-8013
3 |  | Service |  |  |  | GA-110A/B SUCTION |  |  |  |  | FRESH PP TO EA-112 |  |  |  | PPW TO EA-112 |  |  |  | GA-112A/B SUCTION |  |  |  | GA-113 SUCTION |  |  |  |  | ETHER TO WWT |  |  |  | GA-333A/B DISCHARGE |  |  |  | NPA FRACTION TO EA-333 |  |  |  | EA-324 OUTLET |  |  |  |  | GA-353A/B DISCHARGE |  |  |  | FA-354 OUTLET |  |  |  | GA-373A/B DISCHARGE |  |  |  | 0 |  |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  |  |  |  | WT TO FB-835
4 |  | Line Class |  |  |  | FF55C |  |  |  |  | BB59C1 |  |  |  | BB59C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB57C1 |  |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  |  | BB55C1
5 |  | Line Size |  | Line Schedule |  | 1" |  |  |  |  | 2" |  |  |  | 3" |  |  |  | 1" |  |  |  | 1" |  |  |  |  | 1" |  |  |  | 1" |  |  |  | 1" |  |  |  | 3" |  |  |  |  | 2" |  |  |  | 1" |  |  |  | 4" |  |  |  | 2" |  |  |  |  | 6" |  |  |  | 6" |  |  |  | 3" |  |  |  | 3/4" |  |  |  |  | 3"
6 |  | Nace Requirement
7 |  | Enclosure Class |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56
8
9 | SERVICE CONDITION | Fluid |  |  |  | PPA |  |  |  |  | PPF |  |  |  | PPW |  |  |  | PCS |  |  |  | PCS |  |  |  |  | PE |  |  |  | PDI |  |  |  | PDI |  |  |  | WW |  |  |  |  | PIP |  |  |  | PS |  |  |  | PIP |  |  |  | PPR |  |  |  |  | PIP |  |  |  | PIP |  |  |  | PDI |  |  |  | NS |  |  |  |  | WT
10 |  | Flow Rate Unit |  | State |  | m3/h |  |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  |  | L | m3/h |  |  | L
11 | [5 empty] | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR | MIN | MAX | NOR | MIN
12 |  | Flow Rate |  |  |  | 0.0036 |  | 0.003 |  | - | 10.704 | 8.92 |  | - | 23.8692 | 19.891 |  | - | 0.012 | 0.01 |  | - | 0.12 |  | 0.1 |  | - | 0.14308 | 0.11923 |  | - | 1.4232 | 1.186 |  | - | 1.4215 | 1.1846 |  | - | 16.755 |  | 13.9625 |  | - | 10.5768 | 8.814 |  | - | 1.2468 | - |  | - | 60 | 50 |  | - | - |  | - |  | - | - | - |  | - | - | - |  | - | - | - |  | - | - |  | - |  | - | 18.9467 | 15.7891 |  | -
13 |  | Inlet Pressure |  | (kg/cm2g) |  | - |  | 0.01 |  | - | - | 49.15 |  | - | - | 49.75 |  | - | - | 0.71 |  | - | - |  | 0.7 |  | - | - | 3.6 |  | - | - | 6.8 |  | - | - | 5.91 |  | - | - |  | 4.19 |  | - | - | 4.37 |  | - | - | 0.5 |  | - | - | 2.69 |  | - | - |  | - |  | - | - | - |  | - | - | - |  | - | - | - |  | - | - |  | - |  | - | - | 5 |  | -
14 |  | Temperature |  | (deg C) |  | - |  | 20 |  | - | - | 25 |  | - | - | 103 |  | - | - | 20 |  | - | - |  | 20 |  | - | - | 53.09 |  | - | - | 117.31 |  | - | - | 117.31 |  | - | - |  | 40 |  | - | - | 89.94 |  | - | - | 30 |  | - | - | 40 |  | - | - |  | - |  | - | - | - |  | - | - | - |  | - | - | - |  | - | - |  | - |  | - | - | AMB |  | -
15 |  | S.G. @ Op |  | Density @ Op (kg/m3) |  |  |  |  | 1210 |  |  |  | 507.17 |  |  |  | 883.26 |  |  |  | 1220 |  |  |  |  | 1220 |  |  |  | 681.94 |  |  |  | 797.2 |  |  |  | 797.2 |  |  |  |  | 992.5 |  |  |  | 711.52 |  |  |  | 802.08 |  |  |  | 774.8 |  |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  |  | - |  |  |  | 997.27
16 |  | Viscosity @ Op. | (cP) | Molcular Mass |  | 4 |  |  | - |  | 0.0803 |  | - |  | 0.2716 |  | - |  | 5 |  | - |  | 5 |  |  | - |  | 0.2396 |  | - |  | 0.2457 |  | - |  | 0.2457 |  | - |  | 0.653 |  |  | - |  | 0.416 |  | - |  | 0.7273 |  | - |  | 1.34 |  | - |  | - |  |  | - |  | - |  | - |  | - |  | - |  | - |  | - |  | - |  |  | - |  | 0.88997 |  | -
17 |  | Max. Allowable Pressure Drop |  |  | (kg/cm2) | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2
18 |  | Design Pressure |  | (kg/cm2g) |  | 0.5 |  |  |  |  | 55 |  |  |  | 61.1 |  |  |  | 1.58 |  |  |  | 1.58 |  |  |  |  | 5.7 |  |  |  | 8.9 |  |  |  | 8.9 |  |  |  | 10.6 |  |  |  |  | 6.2 |  |  |  | 2 |  |  |  | 6.1 |  |  |  | - |  |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  |  | 7
19 |  | Design Temperature |  | (deg C) |  | 120 |  |  |  |  | 65 |  |  |  | 180 |  |  |  | 120 |  |  |  | 120 |  |  |  |  | 120 |  |  |  | 160 |  |  |  | 160 |  |  |  | 150 |  |  |  |  | 120 |  |  |  | 120 |  |  |  | 120 |  |  |  | - |  |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  |  | AMB
20
21 | METER | Type |  |  |  | CORIOLIS |  |  |  |  | CORIOLIS |  |  |  | CORIOLIS |  |  |  | CORIOLIS |  |  |  | CORIOLIS |  |  |  |  | CORIOLIS |  |  |  | CORIOLIS |  |  |  | CORIOLIS |  |  |  | CORIOLIS |  |  |  |  | CORIOLIS |  |  |  | CORIOLIS |  |  |  | CORIOLIS |  |  |  | CORIOLIS |  |  |  |  | CORIOLIS (NOTE 5) |  |  |  | CORIOLIS (NOTE 5) |  |  |  | CORIOLIS (NOTE 5) |  |  |  | CORIOLIS |  |  |  |  | CORIOLIS
22 |  | Process Connection |  | Size |  | * |  |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  |  | *
23 |  |  |  | Rating (ANSI) |  | 150# RF |  |  |  |  | 600# RF |  |  |  | 600# RF |  |  |  | 150# RF |  |  |  | 150# RF |  |  |  |  | 150# RF |  |  |  | 150# RF |  |  |  | 150# RF |  |  |  | 150# RF |  |  |  |  | 150# RF |  |  |  | 150# RF |  |  |  | 150# RF |  |  |  | 300# RF |  |  |  |  | 150# RF |  |  |  | 150# RF |  |  |  | 150# RF |  |  |  | 150# RF |  |  |  |  | 150# RF
24 |  | Accuracy |  |  |  | ± 0.5% of FS |  |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  | ± 0.5% of FS |  |  |  |  | ± 0.5% of FS
25 |  | Nominal Flow Range |  |  |  | * |  |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  | * |  |  |  |  | *
26 |  |  |  | Case |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD
27 |  | Material |  | Flange |  | 316LSS |  |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  |  | 316LSS
28 |  |  |  | Tube |  | 316LSS |  |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  | 316LSS |  |  |  |  | 316LSS
29
30 | TRANSMITTER | Calibration Flow Range |  |  |  | 0 ~ 1 |  |  |  |  | 0 ~ 15 |  |  |  | 0 ~ 30 |  |  |  | 0 ~ 0.1 |  |  |  | 0 ~ 0.2 |  |  |  |  | 0 ~ 0.2 |  |  |  | 0 ~ 2 |  |  |  | 0 ~ 2 |  |  |  | 0 ~ 20 |  |  |  |  | 0 ~ 15 |  |  |  | 0 ~ 2 |  |  |  | 0 ~ 80 |  |  |  | HOLD |  |  |  |  | HOLD |  |  |  | HOLD |  |  |  | HOLD |  |  |  | HOLD |  |  |  |  | 0 ~ 20
31 |  | Output Signal |  |  |  | 4~20mA HART |  |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  | 4~20mA HART |  |  |  |  | 4~20mA HART
32 |  | Power Supply |  |  |  | AC 110V, 60Hz |  |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  | AC 110V, 60Hz |  |  |  |  | AC 110V, 60Hz
33 |  | Electrical Connection |  |  |  | 1/2" NPT x 2ea |  |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  | 1/2" NPT x 2ea |  |  |  |  | 1/2" NPT x 2ea
34 |  | Mounting Type |  |  |  | INTEGRAL (NOTE3) |  |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  | INTEGRAL (NOTE3) |  |  |  |  | INTEGRAL (NOTE3)
35 |  | Local Display |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD
36
37 | OPTION | Special Cable |  |  |  | NOTE 3 |  |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  |  | NOTE 3
38 |  | Mounting Bracket |  |  |  | NOTE 3 |  |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  | NOTE 3 |  |  |  |  | NOTE 3
39 |  | Batch Controller |  |  |  | YES |  |  |  |  | YES |  |  |  | YES |  |  |  | YES |  |  |  | YES |  |  |  |  | YES |  |  |  | YES |  |  |  | YES |  |  |  | YES |  |  |  |  | YES |  |  |  | YES |  |  |  | YES |  |  |  | YES |  |  |  |  | YES |  |  |  | YES |  |  |  | YES |  |  |  | YES |  |  |  |  | YES
40 |  | Cable Gland |  |  |  | - |  |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  |  | -
41
42 | PROQ. | Manufacturer
43 |  | Model
44
45
[7 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE
[7 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor.
[7 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied.
[7 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature,
[8 empty] | 2" pipe mounting bracket and special cable shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket and special cable shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket and special cable shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket and special cable shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket and special cable shall be supplied by vendor.
[7 empty] | 4. | Flow rate unit shall be converted to kg/hour. | [15 empty] | 4. | Flow rate unit shall be converted to kg/hour. | [15 empty] | 4. | Flow rate unit shall be converted to kg/hour. | [15 empty] | 4. | Flow rate unit shall be converted to kg/hour. | [15 empty] | 4. | Flow rate unit shall be converted to kg/hour.
[58 empty] | 5. | Batch Controller (General Requirement)
- Explosion Proof (KOSHA, KGS, KTL)
- Power Supply : AC 120V, 60Hz
- Function : Totalizer, LCD Indicator, Start / Stop / Reset Button, Local / Remote(DCS)
- Valve Function : 1 Stage Open / 1 Stage Close (Valve Supply by Customer)
- Mounting Type : Self Standing Type with Channel
- Analog Signal (4~20mA) to DCS (Flow Value)
- Discrete Signal to DCS (Run, Stop, Valve Open / Close, Local / Remote Status)
- Control Signal from DCS (Run, Stop)
- Setting Value could be changed by DCS via communication or analog signal (4~20mA).