## Cover
[6 empty] | DATA SHEET | [5 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  |  | : M0556 |  | VORTEX FLOWMETER | [5 empty] | DS-IN-F-006
| PROJECT NAME |  |  | : IPA 50KTA Exp. | [7 empty] | REV. NO.
| LOCATION |  |  | : Yeosu, Korea | [7 empty] | A
IPA
| CLIENT |  | LG Chem Ltd.
| PROJECT |  | IPA 50KTA Expansion
| LOCATION |  | Yeosu, Korea
UNIT
DATA SHEET
FOR
VORTEX FLOWMETER
SERVEONE |  |  |  |  | SERVEONE CONFIDENTIAL INFORMATION: The drawings, designs, ideas, specifications and information contained in this document constitute
confidential information belonging to SERVEONE. Neither this document nor its contents may be used for any purpose or disclosed in any manner
CONSTRUCTION BUSINESS UNIT |  |  |  |  | without the prior written consent of SERVEONE. Reproduction of this document without the prior written consent of SERVEONE or LG Chem.,Ltd.
is prohibited. Any authorized reproduction of this document, or any part thereof, including vendors' shop drawings, shall incorporate this notice."
A |  | MAY.24.'18 |  |  | FOR QUOTAION |  | J.H Jung |  | C.M Lee | K.Y Noh |  |  | H.D Yoon
Rev. |  | Date |  |  | Description |  | Prepared |  | Checked | Approved |  |  | Authorized
JOB NO. |  |  |  |  | DEPARTMENT |  |  | DOCUMENT NO. |  |  | ISSUE DATE
M0556 |  |  |  |  | E&I (INSTRUMENT) |  |  | DS-IN-F-006 |  |  | 2018.05.24

## Vortex Flowmeter
[5 empty] | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT |  |  |  | DATA SHEET | [12 empty] | CLIENT
[18 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem | [16 empty] | LG Chem
[18 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO. | [16 empty] | DOC. NO.
| JOB NO. |  | : M0556 |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006 |  |  |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006 |  |  |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006 |  |  |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006 |  |  |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006 |  |  |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006 |  |  |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006 |  |  |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006 |  |  |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006 |  |  |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006 |  |  |  | VORTEX FLOWMETER | [12 empty] | DS-IN-F-006
| PROJECT NAME |  | : IPA 50KTA Exp. | [14 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO. | [16 empty] | REV. NO.
| LOCATION |  | : Yeosu, Korea | [14 empty] | A | [16 empty] | A | [16 empty] | A | [16 empty] | A | [16 empty] | A | [16 empty] | A | [16 empty] | A | [16 empty] | A | [16 empty] | A | [16 empty] | A | [16 empty] | A
1 | GENERAL | Tag No. |  |  | FT-1136 |  |  |  |  | FT-1143 |  |  |  | FT-1145 |  |  |  | FT-1157 |  |  |  | FT-1163 |  |  |  |  | FT-1210 |  |  |  | FT-1211 |  |  |  | FT-1212 |  |  |  | FT-1213 |  |  |  |  | FT-1220 |  |  |  | FT-1221 |  |  |  | FT-1225 |  |  |  | FT-1226 |  |  |  |  | FT-1227 |  |  |  | FT-1312 |  |  |  | FT-1315 |  |  |  | FT-1316 |  |  |  |  | FT-1325 |  |  |  | FT-1330 |  |  |  | FT-1340 |  |  |  | FT-1341 |  |  |  |  | FT-1353 |  |  |  | FT-1356 |  |  |  | FT-1357 |  |  |  | FT-1358 |  |  |  |  | FT-1359 |  |  |  | FT-1361 |  |  |  | FT-1366 |  |  |  | FT-1368 |  |  |  |  | FT-1388 |  |  |  | FT-1430 |  |  |  | FT-1802 |  |  |  | FT-1803 |  |  |  |  | FT-1804 |  |  |  | FT-1818 |  |  |  | FT-1820 |  |  |  | FT-1830 |  |  |  |  | FT-1836 |  |  |  | FT-1838 |  |  |  | FT-1842 |  |  |  | FT-1847 |  |  |  |  | FT-1870 |  |  |  | FT-1880 |  |  |  | FT-1890 |  |  |  | 44
2 |  | P&ID No. |  |  | D-1-IPA-1001 |  |  |  |  | D-1-IPA-1002 |  |  |  | D-1-IPA-1002 |  |  |  | D-1-IPA-1003 |  |  |  | D-1-IPA-1003 |  |  |  |  | D-1-IPA-2001 |  |  |  | D-1-IPA-2001 |  |  |  | D-1-IPA-2001 |  |  |  | D-1-IPA-2001 |  |  |  |  | D-1-IPA-2002 |  |  |  | D-1-IPA-2002 |  |  |  | D-1-IPA-2002 |  |  |  | D-1-IPA-2002 |  |  |  |  | D-1-IPA-2002 |  |  |  | D-1-IPA-3001 |  |  |  | D-1-IPA-3001 |  |  |  | D-1-IPA-3001 |  |  |  |  | D-1-IPA-3003 |  |  |  | D-1-IPA-3002 |  |  |  | D-1-IPA-3003 |  |  |  | D-1-IPA-3003 |  |  |  |  | D-1-IPA-3004 |  |  |  | D-1-IPA-3004 |  |  |  | D-1-IPA-3004 |  |  |  | D-1-IPA-3004 |  |  |  |  | D-1-IPA-3007 |  |  |  | D-1-IPA-3006 |  |  |  | D-1-IPA-3006 |  |  |  | D-1-IPA-3006 |  |  |  |  | D-1-IPA-3003 |  |  |  | D-1-IPA-4003 |  |  |  | D-1-IPA-8011 |  |  |  | D-1-IPA-8011 |  |  |  |  | D-1-IPA-8011 |  |  |  | D-1-IPA-8011 |  |  |  | D-1-IPA-8012 |  |  |  | D-1-IPA-8013 |  |  |  |  | D-1-IPA-8013 |  |  |  | D-1-IPA-8013 |  |  |  | D-1-IPA-8014 |  |  |  | D-1-IPA-8014 |  |  |  |  | D-1-IPA-8016 |  |  |  | D-1-IPA-8022 |  |  |  | D-1-IPA-8024
3 |  | Service |  |  | PPW TO EA-112 |  |  |  |  | GAS BLEED TO DA-210 |  |  |  | WASH WATER TO DA-140 |  |  |  | FC-150 GAS OUTLET |  |  |  | GA-163A/B DISCHARGE |  |  |  |  | DA-210 BOTTOMS |  |  |  | SH TO EA-211 |  |  |  | DA-210 REFLUX |  |  |  | DA-220 FEED INLET |  |  |  |  | DA-220 BOTTOMS |  |  |  | SL TO EA-221 |  |  |  | DA-220 REFLUX |  |  |  | EA-225 VENT GAS |  |  |  |  | GA-226A/B DISCHARGE |  |  |  | DA-310 FEED INLET |  |  |  | DA-310 REFLUX |  |  |  | ETHER TO EA-112 |  |  |  |  | DA-320 REFLUX |  |  |  | DA-330 FEED INLET |  |  |  | ACETONE TO FA-781B |  |  |  | DA-320 OVHD TO DA-340 |  |  |  |  | EA-353A/B OUTLET |  |  |  | DA-360 FEED INLET |  |  |  | DA-350 REFLUX |  |  |  | DA-350 FEED INLET |  |  |  |  | GA-371B DISCHARGE |  |  |  | SL TO EA-361 |  |  |  | GA-365A/B DISCHARGE |  |  |  | DA-360 REFLUX |  |  |  |  | WT TO EA-323A/B |  |  |  | 0 |  |  |  | BFW FORM B.L |  |  |  | CL TO B.L |  |  |  |  | CL TO C.W BASIN |  |  |  | CL TO BG-802 |  |  |  | C.W TO BLOW DOWN |  |  |  | SM TO DA-830 |  |  |  |  | WT TO DA-830 |  |  |  | GA-830A/B DISCHARGE |  |  |  | AI FROM BL |  |  |  | NL FROM BL |  |  |  |  | 0 |  |  |  | 0 |  |  |  | WW FROM EA-324
4 |  | Line Class |  |  | BB59C1 |  |  |  |  | BB59C1 |  |  |  | BB59C1 |  |  |  | BB57C1 |  |  |  | BB55C1 |  |  |  |  | BB57C1 |  |  |  | BA59C |  |  |  | BB57C1 |  |  |  | BB57C1 |  |  |  |  | BB57C1 |  |  |  | BB55C |  |  |  | BB59C1 |  |  |  | BB55C1 |  |  |  |  | BB59C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB59C1 |  |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  |  | BB55C1 |  |  |  | BB55C |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | A1A |  |  |  | BB55C |  |  |  |  | BB55C |  |  |  | BB55C |  |  |  | BB55C |  |  |  | BB57C |  |  |  |  | BB55C1 |  |  |  | BB55C1 |  |  |  | BC55C |  |  |  | BB55C |  |  |  |  | BB55C |  |  |  | BB55C |  |  |  | BB55C1
5 |  | Line Size |  | Line Schedule | 3" |  |  |  |  | 1" |  |  |  | 3" |  |  |  | 3" |  |  |  | 1 1/2" |  |  |  |  | 1 1/2" |  |  |  | 1 1/2" |  |  |  | 2" |  |  |  | 2" |  |  |  |  | 1" |  |  |  | 2" |  |  |  | 2" |  |  |  | 1" |  |  |  |  | 2" |  |  |  | 4" |  |  |  | 1" |  |  |  | 1" |  |  |  |  | 3" |  |  |  | 1 1/2" |  |  |  | 1" |  |  |  | 4" |  |  |  |  | 2" |  |  |  | 1" |  |  |  | 4" |  |  |  | 2" |  |  |  |  | 1 1/2" |  |  |  | 4" |  |  |  | 1" |  |  |  | 1" |  |  |  |  | 3" |  |  |  | 1" |  |  |  | 1 1/2" |  |  |  | 4" |  |  |  |  | 4" |  |  |  | 1 1/2" |  |  |  | 4" |  |  |  | 4" |  |  |  |  | 3" |  |  |  | 3" |  |  |  | 4" |  |  |  | 3" |  |  |  |  | 3" |  |  |  | 4" |  |  |  | 2"
6 |  | Nace Requirement
7 |  | Enclosure Class |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56 |  |  |  | Ex d IIB T4, IP 56
8
9 | SERVICE CONDITION | Fluid |  |  | PPW |  |  |  |  | PRG |  |  |  | PWW |  |  |  | PRG |  |  |  | PDI |  |  |  |  | PHE |  |  |  | SH |  |  |  | PRG |  |  |  | PRG |  |  |  |  | PHE |  |  |  | SL |  |  |  | PPR |  |  |  | PRG |  |  |  |  | PPR |  |  |  | PDI |  |  |  | PDI |  |  |  | PE |  |  |  |  | PDI |  |  |  | PDI |  |  |  | PDI |  |  |  | PDI |  |  |  |  | PIP |  |  |  | PDI |  |  |  | PS |  |  |  | PDI |  |  |  |  | PDI |  |  |  | SL |  |  |  | PDI |  |  |  | PDI |  |  |  |  | WT |  |  |  | WT |  |  |  | BFW |  |  |  | CL |  |  |  |  | CL |  |  |  | CL |  |  |  | WCS |  |  |  | SM |  |  |  |  | WT |  |  |  | WT |  |  |  | AI |  |  |  | NL |  |  |  |  | AP |  |  |  | WR |  |  |  | WW
10 |  | Flow Rate Unit |  | State | m3/h |  |  |  | L | kg/h |  |  | V | m3/h |  |  | L | kg/h |  |  | V | m3/h |  |  |  | L | m3/h |  |  | L | kg/h |  |  | S | m3/h |  |  | L | m3/h |  |  |  | L | m3/h |  |  | L | kg/h |  |  | S | m3/h |  |  | L | kg/h |  |  |  | V | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  | L | kg/h |  |  |  | V | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  |  | L | m3/h |  |  | L | kg/h |  |  | S | m3/h |  |  | L | m3/h |  |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  |  | L | m3/h |  |  | L | m3/h |  |  | L | m3/h |  |  | L | kg/h |  |  |  | S | m3/h |  |  | L | m3/h |  |  | L | Nm3/h |  |  | V | Nm3/h |  |  |  | V | kg/h |  |  | V | m3/h |  |  | L | m3/h |  |  | L
11 |  |  |  |  | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX |  | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN | MAX | NOR |  | MIN
12 |  | Flow Rate |  |  | 23.8692 |  | 19.891 |  | - | 2000.88 | 1667.4 |  | - | 39.1968 | 32.664 |  | - | 2253.24 | 1877.7 |  | - | 1.752 |  | 1.46 |  | - | 1.4724 | 1.227 |  | - | 333.36 | 277.8 |  | - | 10.7776 | 8.9813 |  | - | 10.1241 |  | 8.4368 |  | - | 0.07666 | 0.063884 |  | - | 1094.4 | 912 |  | - | 7.1242 | 5.9369 |  | - | 32.4 |  | 27 |  | - | 10.38 | 8.65 |  | - | 71.0568 | 59.214 |  | - | 1.2612 | 1.051 |  | - | 1.218 |  | 1.015 |  | - | 32.3148 | 26.929 |  | - | 7.2912 | 6.076 |  | - | 0.020548 | 0.017123 |  | - | 900 |  | 750 |  | - | 9.7116 | 8.093 |  | - | 1.86 | 1.55 |  | - | 39.4476 | 32.873 |  | - | 12.0204 |  | 10.017 |  | - | 7.9995 | 6.6662 |  | - | 849.6 | 708 |  | - | 0.8484 | 0.707 |  | - | 1.1892 |  | 0.991 |  | - | HOLD | HOLD |  | - | - | - |  | - | 2747 | 1948 |  | - | HOLD |  | HOLD |  | - | HOLD | HOLD |  | - | HOLD | HOLD |  | - | HOLD | HOLD |  | - | 1154.4 |  | 962 |  | - | 20 | 16 |  | - | 19.7112 | 16.426 |  | - | 1000 | 833 |  | - | 172 |  | 84 |  | - | - | - |  | - | - | - |  | - | 16.755 | 13.9625 |  | -
13 |  | Inlet Pressure |  | (kg/cm2g) | - |  | 49.55 |  | - | - | 44.8 |  | - | - | 39.25 |  | - | - | 16.8 |  | - | - |  | 5.7 |  | - | - | 16.53 |  | - | - | 31 |  | - | - | 22.9 |  | - | - |  | 22.9 |  | - | - | 19.59 |  | - | - | 3.5 |  | - | - | 22.07 |  | - | - |  | 19 |  | - | - | 50.92 |  | - | - | 4.98 |  | - | - | 5.2 |  | - | - |  | 50.8 |  | - | - | 7.97 |  | - | - | 6.37 |  | - | - | 0.5 |  | - | - |  | 1.6 |  | - | - | 3.29 |  | - | - | 3.44 |  | - | - | 4.7 |  | - | - |  | 1.67 |  | - | - | 5.19 |  | - | - | 3.5 |  | - | - | 4.14 |  | - | - |  | 4.14 |  | - | - | 5.58 |  | - | - | - |  | - | - | - |  | - | - |  | - |  | - | - | - |  | - | - | - |  | - | - | 5 |  | - | - |  | 9.8 |  | - | - | 5.35 |  | - | - | 7.3 |  | - | - | 7 |  | - | - |  | 7 |  | - | - | - |  | - | - | - |  | - | - | 7 |  | -
14 |  | Temperature |  | (deg C) | - |  | 102.87 |  | - | - | 120.12 |  | - | - | 95 |  | - | - | 118.2 |  | - | - |  | 77.7 |  | - | - | 172.91 |  | - | - | 255 |  | - | - | 39.8 |  | - | - |  | 39.8 |  | - | - | 54.09 |  | - | - | 147 |  | - | - | 48.3 |  | - | - |  | 47.19 |  | - | - | 47 |  | - | - | 82.39 |  | - | - | 45 |  | - | - |  | 53.09 |  | - | - | 50.19 |  | - | - | 111.3 |  | - | - | 56 |  | - | - |  | 105.62 |  | - | - | 40 |  | - | - | 40.3 |  | - | - | 40 |  | - | - |  | 106.38 |  | - | - | 40 |  | - | - | 147 |  | - | - | 74.07 |  | - | - |  | 74.07 |  | - | - | AMB |  | - | - | - |  | - | - | HOLD |  | - | - |  | HOLD |  | - | - | HOLD |  | - | - | HOLD |  | - | - | 32 |  | - | - |  | 198 |  | - | - | 100 |  | - | - | 20 |  | - | - | AMB |  | - | - |  | AMB |  | - | - | - |  | - | - | - |  | - | - | 40 |  | -
15 |  | S.G. @ Op |  | Density @ Op |  |  |  | 883.26 |  |  |  | 90.02 |  |  |  | 963 |  |  |  | 27.8 |  |  |  |  | 739.12 |  |  |  | 545.28 |  |  |  | 14.2 |  |  |  | 454.4 |  |  |  |  | 454.4 |  |  |  | 442.21 |  |  |  | 1.87 |  |  |  | 465.7 |  |  |  |  | 43.43 |  |  |  | 436.8 |  |  |  | 880.7 |  |  |  | 969.38 |  |  |  |  | 681.94 |  |  |  | 787.17 |  |  |  | 738.89 |  |  |  | 795.41 |  |  |  |  | 3.98 |  |  |  | 774.84 |  |  |  | 906.48 |  |  |  | 802.08 |  |  |  |  | 712.8 |  |  |  | 774.8 |  |  |  | 2.34 |  |  |  | 764.45 |  |  |  |  | 764.45 |  |  |  | 998.5 |  |  |  | - |  |  |  | HOLD |  |  |  |  | HOLD |  |  |  | HOLD |  |  |  | HOLD |  |  |  | 995.3 |  |  |  |  | 5.0955 |  |  |  | 958.6 |  |  |  | 958.6 |  |  |  | 8.08 |  |  |  |  | 6.367 |  |  |  | - |  |  |  | - |  |  |  | 992.5
16 |  | Viscosity @ Op. |  | Molcular Mass | 0.2716 |  |  | - |  | 0.0143 |  | 43.87 |  | 0.296 |  | - |  | 0.011 |  | 43.78 |  | 0.3806 |  |  | - |  | 0.125 |  | - |  | 0.019 |  | 18 |  | 0.0657 |  | - |  | 0.0657 |  |  | - |  | 0.0725 |  | - |  | 0.014 |  | 18 |  | 0.076 |  | - |  | 0.0105 |  |  | 41.89 |  | 0.056 |  | - |  | 0.32 |  | - |  | 0.611 |  | - |  | 0.2396 |  |  | - |  | 0.8481 |  | - |  | 0.2658 |  | - |  | 0.727 |  | - |  | 0.011 |  |  | 46.92 |  | 1.3404 |  | - |  | 0.7083 |  | - |  | 0.7273 |  | - |  | 0.2885 |  |  | - |  | 1.34 |  | - |  | 0.014 |  | 18 |  | 0.456 |  | - |  | 0.456 |  |  | - |  | 1.0014 |  | - |  | - |  | - |  | HOLD |  | - |  | HOLD |  |  | - |  | HOLD |  | - |  | HOLD |  | - |  | 0.77 |  | - |  | 0.016 |  |  | 18 |  | 0.28 |  | - |  | 0.28 |  | - |  | 0.018566 |  | 28.95 |  | 0.0178 |  |  | 28 |  | - |  | - |  | - |  | - |  | 0.653 |  | -
17 |  | Max. Allowable Pressure Drop (kg/cm2g) |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  |  | 0.2 |  |  |  | 0.2 |  |  |  | 0.2
18 |  | Design Pressure |  | (kg/cm2g) | 61.1 |  |  |  |  | 55 |  |  |  | 48.1 |  |  |  | 20 |  |  |  | 7.6 |  |  |  |  | 25 |  |  |  | 46 |  |  |  | 31.7 |  |  |  | 31.7 |  |  |  |  | 25 |  |  |  | 6 |  |  |  | 28.3 |  |  |  | 25 |  |  |  |  | 55 |  |  |  | 7.9 |  |  |  | 7.4 |  |  |  | 63.5 |  |  |  |  | 10.2 |  |  |  | 8.5 |  |  |  | 2 |  |  |  | 4 |  |  |  |  | 6.2 |  |  |  | 5.7 |  |  |  | 7.1 |  |  |  | 4 |  |  |  |  | 6.7 |  |  |  | 6 |  |  |  | 6.9 |  |  |  | 6.9 |  |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  | - |  |  |  |  | - |  |  |  | - |  |  |  | 7 |  |  |  | 12 |  |  |  |  | 10 |  |  |  | 10.5 |  |  |  | 10 |  |  |  | 10 |  |  |  |  | - |  |  |  | - |  |  |  | 10.6
19 |  | Design Temperature |  | (deg C) | 180 |  |  |  |  | 180 |  |  |  | 150 |  |  |  | 140 |  |  |  | 120 |  |  |  |  | 200 |  |  |  | 400 |  |  |  | 120 |  |  |  | 120 |  |  |  |  | 120 |  |  |  | 200 |  |  |  | 120 |  |  |  | 120 |  |  |  |  | 120 |  |  |  | 120 |  |  |  | 120 |  |  |  | 120 |  |  |  |  | 140 |  |  |  | 180 |  |  |  | 82 |  |  |  | 180 |  |  |  |  | 120 |  |  |  | 120 |  |  |  | 120 |  |  |  | 180 |  |  |  |  | 120 |  |  |  | 200 |  |  |  | 120 |  |  |  | 120 |  |  |  |  | AMB |  |  |  | - |  |  |  | HOLD |  |  |  | HOLD |  |  |  |  | HOLD |  |  |  | HOLD |  |  |  | 50 |  |  |  | 250 |  |  |  |  | 160 |  |  |  | HOLD |  |  |  | 60 |  |  |  | 60 |  |  |  |  | - |  |  |  | - |  |  |  | 150
20
21 | METER | Connection Size |  | Rating (ANSI) | * |  |  | 600# RF |  | * |  | 600# RF |  | * |  | 600# RF |  | * |  | 300# RF |  | * |  |  | 150# RF |  | * |  | 300# RF |  | * |  | 600# RF |  | * |  | 300# RF |  | * |  |  | 300# RF |  | * |  | 300# RF |  | * |  | 150# RF |  | * |  | 600# RF |  | * |  |  | 150# RF |  | * |  | 600# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  |  | 600# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  | 600# RF |  | * |  |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  |  | 300# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF |  | * |  | 150# RF
22 |  | Sensor Type |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD
23 |  | Accuracy |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD |  |  |  | ± 1% OF RD
24 |  | Calibrated Flow Range |  |  | 0 ~ 30 |  |  |  |  | 0 ~ 2500 |  |  |  | 0 ~ 50 |  |  |  | 0 ~ 2500 |  |  |  | 0 ~ 3 |  |  |  |  | 0 ~ 2 |  |  |  | 0 ~ 400 |  |  |  | 0 ~ 15 |  |  |  | 0 ~ 15 |  |  |  |  | 0 ~ 0.1 |  |  |  | 0 ~ 1500 |  |  |  | 0 ~ 10 |  |  |  | 0 ~ 40 |  |  |  |  | 0 ~ 15 |  |  |  | 0 ~ 80 |  |  |  | 0 ~ 2 |  |  |  | 0 ~ 2 |  |  |  |  | 0 ~ 40 |  |  |  | 0 ~ 10 |  |  |  | 0 ~ 0.05 |  |  |  | 0 ~ 1000 |  |  |  |  | 0 ~ 15 |  |  |  | 0 ~ 3 |  |  |  | 0 ~ 50 |  |  |  | 0 ~ 15 |  |  |  |  | 0 ~ 15 |  |  |  | 0 ~ 1000 |  |  |  | 0 ~ 1.5 |  |  |  | 0 ~ 2 |  |  |  |  | HOLD |  |  |  | HOLD |  |  |  | 0 ~ 3000 |  |  |  | HOLD |  |  |  |  | HOLD |  |  |  | HOLD |  |  |  | HOLD |  |  |  | 0 ~ 1500 |  |  |  |  | 0 ~ 25 |  |  |  | 0 ~ 25 |  |  |  | 0 ~ 1500 |  |  |  | 0 ~ 150 |  |  |  |  | HOLD |  |  |  | HOLD |  |  |  | 0 ~ 20
25 |  |  |  | Case | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  |  | MFR. STD |  |  |  | MFR. STD |  |  |  | MFR. STD
26 |  | Material |  | Body | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS
27 |  |  |  | Wetted Part | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS |  |  |  |  | 316SS |  |  |  | 316SS |  |  |  | 316SS
28
29
30 | TRANSMITTER | Scale Range
31 |  | Mounting Type |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3) |  |  |  | MFR. STD (NOTE 3)
32 |  | Output Signal |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART |  |  |  | 4 ~ 20 mA HART
33 |  | Power Supply |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  |  | LOOP POWER |  |  |  | LOOP POWER |  |  |  | LOOP POWER
34 |  | Electrical Connection |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT |  |  |  | 1/2" NPT
35
36
37 | ACCESSORIES | Display |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD |  |  |  |  | LCD |  |  |  | LCD |  |  |  | LCD
38 |  | Center Ring
39 |  | Mounting Bracket |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  |  | (NOTE 3) |  |  |  | (NOTE 3) |  |  |  | (NOTE 3)
40 |  | Cable Gland
41 |  | Totalizer
42
43 | PROQ. | Manufacturer
44 |  | Model
45
46
[6 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE | [16 empty] | NOTE
[6 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor. | [15 empty] | 1. | " * " Marked Items shall be specified or recommended by Vendor.
[6 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied. | [15 empty] | 2. | Stainless steel name plate shall be supplied.
[6 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature, | [15 empty] | 3. | Generaly intergral type shall be used. When remote type is used because of process temperature,
[7 empty] | 2" pipe mounting bracket shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket shall be supplied by vendor. | [16 empty] | 2" pipe mounting bracket shall be supplied by vendor.