## Cover
[6 empty] | DATA SHEET | [5 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  |  | : M0556 |  | LOCAL INSTRUMENT (PRESSURE GAUGE) | [5 empty] | DS-IN-F-018A
| PROJECT NAME |  |  | : IPA 50KTA Exp. | [7 empty] | REV. NO.
| LOCATION |  |  | : Yeosu, Korea | [7 empty] | A
IPA
| CLIENT |  | LG Chem Ltd.
| PROJECT |  | IPA 50KTA Expansion
| LOCATION |  | Yeosu, Korea
UNIT
DATA SHEET
FOR
LOCAL INSTRUMENT
(PRESSURE GAUGE)
SERVEONE |  |  |  |  | SERVEONE CONFIDENTIAL INFORMATION: The drawings, designs, ideas, specifications and information contained in this document constitute
confidential information belonging to SERVEONE. Neither this document nor its contents may be used for any purpose or disclosed in any manner
CONSTRUCTION BUSINESS UNIT |  |  |  |  | without the prior written consent of SERVEONE. Reproduction of this document without the prior written consent of SERVEONE or LG Chem.,Ltd.
is prohibited. Any authorized reproduction of this document, or any part thereof, including vendors' shop drawings, shall incorporate this notice."
A |  | AUG.16.'18 |  |  | FOR QUOTAION |  | J.H.Jung |  | C.M Lee | K.Y Noh |  |  | H.D.Yoon
Rev. |  | Date |  |  | Description |  | Prepared |  | Checked | Approved |  |  | Authorized
JOB NO. |  |  |  |  | DEPARTMENT |  |  | DOCUMENT NO. |  |  | ISSUE DATE
M0556 |  |  |  |  | E&I (INSTRUMENT) |  |  | DS-IN-F-018A |  |  | 2018.08.16

## Flanged_PG_150
[9 empty] | DATA SHEET | [13 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  |  | : M0556 |  |  |  |  | LOCAL INSTRUMENT (PRESSURE GAUGE) | [13 empty] | DS-IN-F-018A
| PROJECT NAME |  |  | : IPA 50KTA Exp. | [18 empty] | REV. NO.
| LOCATION |  |  | : Yeosu, Korea | [18 empty] | A
GENERAL | [12 empty] | ELEMENT
1. | TYPE |  | ■ | DIRECT READING |  |  |  | □ | RECEIVER |  |  |  | 13. | ELEMENT TYPE |  |  |  | ■ | BOURDON |  |  |  | □ | BELLOWS
2. | MOUNTING |  | ■ | LOCAL |  | □ | SURFACE |  |  | □ | FLUSH |  | 14. | ELEMENT MAT'L |  |  |  | ■ | 316SS |  | □ | OTHER
3. | DIAL SIZE |  | ■ | 100 mm |  | □ | 150 mm |  |  | □ | 4 1/2" |  | 15. | SOCKET MAT'L |  |  |  | ■ | 316SS |  | □ | OTHER
4. | DIAL COLOR |  | ■ | WHITE BACK GROUND BLACK LETTERS | [8 empty] | 16. | CONNECTION |  |  |  | ■ | 1/2" NPT(M) |  | □ | OTHER
5. | CASE MAT'L |  | ■ | 304SS |  | □ | 316SS |  |  | □ | ALUMINUM |  | 17. | MOVEMENT |  |  |  | ■ | 304SS |  | □ | OTHER
6. | RING TYPE |  | □ | SCREWED |  | ■ | BAYONET |  |  | □ | SLIP |  | 18. | DIAPHRAGM SEAL
7. | ACCURACY |  | ■ | ± 1.0 % OF F.S |  |  | □ | OTHER | [6 empty] | DIAPHRAGM MAT'L |  |  |  |  | 316LSS
8. | OVERRANGE PROTECTION |  |  |  | ■ | 1.3 x F.S | [8 empty] | UPPER FLANGE MAT'L |  |  |  |  | 316SS
9. | BLOW OUT PROTECTION |  |  |  | ■ | BACK |  |  |  | □ | SIDE |  |  |  | UNDER FLANGE MAT'L |  |  |  |  | 316SS
[5 empty] | ■ | SOLID FRONT ( ≥ 30 kg/cm2g) | [8 empty] | FILL FLUID |  |  |  |  | SILICONE OIL
10. | LENS |  | ■ | SAFETY GLASS |  |  | □ | OTHER | [6 empty] | PROCESS CONNECTION |  |  |  |  | 1" 150# RF
OPTION | [12 empty] | 19. | MFR & MODEL NO.
11. | ACCESSORIES |  | □ | SYPHONE (S) |  |  |  | □ | SNUBBER (B)
| (SEE LIST) |  | ■ | DAMPENER (D) |  |  |  | □ | SOLID FRONT (SF) |  |  |  | NOTE
|  |  | ■ | LIQUID FILLED (LF) | [8 empty] | 1. | Syphone Tube shall be min. 80 sch.
12. | MANDATORY OPTION | [11 empty] | 2. | Vendor shall be take consideration the service condition for seal fluid
|  |  | □ | GAUGE UNION (G) |  |  |  | □ | MANIFOLD VALVE (M) |  |  |  |  | and accessories.
|  |  | ■ | COOLING FIN (C) - MFR. STD.
NO | TAG NO | SERVICE |  |  |  |  | P&ID NO |  |  | FLUID |  | STATE | RANGE (kg/cm2g) |  |  | PRESSURE (kg/cm2g) |  |  | TEMP. (deg C) |  |  | ACCESSORIES |  |  | REMARK
[29 empty] | DESIGN PRESS. |  | DESIGN TEMP. |  | REMARK
[16 empty] | MAX | NOR |  | MAX | NOR | [5 empty] | TYPE | SPEC. | RATING | INT. | EXT. | HOT | COLD
1 | PG-1134 | FA-134 EQUIP. |  |  |  |  | D-1-IPA-0004-2 |  |  | PRG |  | L | -1 ~ 1 |  |  |  | 0.5 |  |  | 25 |  |  |  | LF |  | 1" 150# RF | FF55C | 150# RF | 8 | 0 | 60 | 0 |  | O
2 | PG-1135 | FA-134 EQUIP. |  |  |  |  | D-1-IPA-0004-2 |  |  | PRG |  | L | -1 ~ 1 |  |  |  | 0.5 |  |  | 25 |  |  |  | LF |  | 1" 150# RF | FF55C | 150# RF | 8 | 0 | 60 | 0 |  | O
3 | PG-1136 | GA-134 DISCH. |  |  |  |  | D-1-IPA-0004-2 |  |  | PDI |  | L | 0 ~ 5 |  |  |  | 2.47 |  |  | 25 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10.7 | 0 | 60 | 0 |  | O
4 | PG-1137 | FA-134 EQUIP. |  |  |  |  | D-1-IPA-0004-2 |  |  | PDI |  | L | -1 ~ 1 |  |  |  | 0.5 |  |  | 25 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 8 | 0 | 60 | 0 |  | O
5 | PG-1138 | GA-112A/B SUCTION |  |  |  |  | D-1-IPA-1001 |  |  | PCS |  | L | -1 ~ 1 |  |  | 1.58 | ATM |  |  | 20 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 0.5 | 0 | 120 | 0 |  | O
6 | PG-1155A | GA-155A DISCH. |  |  |  |  | D-1-IPA-1003 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.39 |  |  | 82 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.9 | 0 | 120 | 0 |  | O
7 | PG-1155B | GA-155B DISCH. |  |  |  |  | D-1-IPA-1003 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.39 |  |  | 82 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.9 | 0 | 120 | 0 |  | O
8 | PG-1163A | GA-163A DISCH. |  |  |  |  | D-1-IPA-1003 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.71 |  |  | 78 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.6 | 0 | 120 | 0 |  | O
9 | PG-1163B | GA-163B DISCH. |  |  |  |  | D-1-IPA-1003 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.71 |  |  | 78 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.6 | 0 | 120 | 0 |  | O
10 | PG-1170-1 | FC-170 EQUIP. |  |  |  |  | D-1-IPA-1003 |  |  | PRG |  | V | -1 ~ 1 |  |  |  | 0.1 |  |  | 38 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 2 | 0 | 120 | 0 |  | O
11 | PG-1221 | LP TO EA-221 |  |  |  |  | D-1-IPA-2002 |  |  | SL |  | S | 0 ~ 5 |  |  |  | 2.5 |  |  | 138.2 |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 6 | 0 | 200 | 0 |  | O
12 | PG-1313A | GA-313A DISCH. |  |  |  |  | D-1-IPA-3001 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.73 |  |  | 97 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.8 | 0 | 150 | 0 |  | O
13 | PG-1313B | GA-313B DISCH. |  |  |  |  | D-1-IPA-3001 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.73 |  |  | 97 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.8 | 0 | 150 | 0 |  | O
14 | PG-1315A | GA-315A DISCH. |  |  |  |  | D-1-IPA-3001 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.25 |  |  | 45 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.4 | 0 | 120 | 0 |  | O
15 | PG-1315B | GA-315B DISCH. |  |  |  |  | D-1-IPA-3001 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.25 |  |  | 45 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.4 | 0 | 120 | 0 |  | O
16 | PG-1317A | GA-317A DISCH. |  |  |  |  | D-1-IPA-3001 |  |  | PE |  | L | 0 ~ 10 |  |  |  | 3.57 |  |  | 45 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.7 | 0 | 120 | 0 |  | O
17 | PG-1317B | GA-317B DISCH. |  |  |  |  | D-1-IPA-3001 |  |  | PE |  | L | 0 ~ 10 |  |  |  | 3.57 |  |  | 45 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.7 | 0 | 120 | 0 |  | O
18 | PG-1318 | N2 TO FLARE |  |  |  |  | D-1-IPA-3001 |  |  | NL |  | V | -1 ~ 1 |  |  |  | 0.01 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
19 | PG-1319 | DA-310 EQUIP. |  |  |  |  | D-1-IPA-3001 |  |  | PDI |  | L | -1 ~ 1 |  |  |  | 0.25 |  |  | 97 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 150 | 0 |  | O
20 | PG-1323A | GA-323A DISCH. |  |  |  |  | D-1-IPA-3002 |  |  | PWW |  | L | 0 ~ 15 |  |  |  | 8.23 |  |  | 133 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10.6 | 0 | 180 | 0 |  | O
21 | PG-1323B | GA-323B DISCH. |  |  |  |  | D-1-IPA-3002 |  |  | PWW |  | L | 0 ~ 15 |  |  |  | 8.23 |  |  | 133 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10.6 | 0 | 180 | 0 |  | O
22 | PG-1325A | GA-325A DISCH. |  |  |  |  | D-1-IPA-3003 |  |  | PDI |  | L | 0 ~ 15 |  |  |  | 7.97 |  |  | 50 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10.2 | 0 | 140 | 0 |  | O
23 | PG-1325B | GA-325B DISCH. |  |  |  |  | D-1-IPA-3003 |  |  | PDI |  | L | 0 ~ 15 |  |  |  | 7.97 |  |  | 50 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10.2 | 0 | 140 | 0 |  | O
24 | PG-1327 | N2 TO FLARE |  |  |  |  | D-1-IPA-3003 |  |  | NL |  | V | -1 ~ 1 |  |  |  | 0.01 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
25 | PG-1328 | FA-325 EQUIP. |  |  |  |  | D-1-IPA-3003 |  |  | PDI |  | L | -1 ~ 1 |  |  |  | 0.02 |  |  | 50 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 140 | 0 |  | O
26 | PG-1330A | GA-330A DISCH. |  |  |  |  | D-1-IPA-3002 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 6.37 |  |  | 111 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.5 | 0 | 180 | 0 |  | O
27 | PG-1330B | GA-330B DISCH. |  |  |  |  | D-1-IPA-3002 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 6.37 |  |  | 111 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.5 | 0 | 180 | 0 |  | O
28 | PG-1331 | LP TO EA-331 |  |  |  |  | D-1-IPA-3002 |  |  | SL |  | S | 0 ~ 5 |  |  |  | 2.5 |  |  | 138.2 |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 6 | 0 | 200 | 0 |  | O
29 | PG-1331A | GA-333A DISCH. |  |  |  |  | D-1-IPA-3002 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 6.8 |  |  | 117 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.9 | 0 | 160 | 0 |  | O
30 | PG-1331B | GA-333B DISCH. |  |  |  |  | D-1-IPA-3002 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 6.8 |  |  | 117 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.9 | 0 | 160 | 0 |  | O
31 | PG-1345 | N2 TO FLARE |  |  |  |  | D-1-IPA-3003 |  |  | NL |  | V | -1 ~ 1 |  |  |  | 0.005 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
32 | PG-1351 | DA-350 EQUIP. |  |  |  |  | D-1-IPA-3004 |  |  | PIP |  | L | -1 ~ 2 |  |  |  | 0.4 |  |  | 90 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 120 | 0 |  | O
33 | PG-1351A | EA-351A EQUIP. |  |  |  |  | D-1-IPA-3004 |  |  | PDI |  | V | 0 ~ 5 |  |  |  | 1.6 |  |  | 105.6 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 4 | 0 | 150 | 0 |  | O
34 | PG-1351B | EA-351B EQUIP. |  |  |  |  | D-1-IPA-3004 |  |  | PDI |  | V | 0 ~ 5 |  |  |  | 1.6 |  |  | 105.6 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 4 | 0 | 150 | 0 |  | O
35 | PG-1353A | GA-353A DISCH. |  |  |  |  | D-1-IPA-3004 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.37 |  |  | 90 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6 | 0 | 120 | 0 |  | O
36 | PG-1353B | GA-353B DISCH. |  |  |  |  | D-1-IPA-3004 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.37 |  |  | 90 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6 | 0 | 120 | 0 |  | O
37 | PG-1355A | GA-355A DISCH. |  |  |  |  | D-1-IPA-3004 |  |  | PS |  | L | 0 ~ 10 |  |  |  | 4.73 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.1 | 0 | 120 | 0 |  | O
38 | PG-1355B | GA-355B DISCH. |  |  |  |  | D-1-IPA-3004 |  |  | PS |  | L | 0 ~ 10 |  |  |  | 4.73 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.1 | 0 | 120 | 0 |  | O
39 | PG-1356A | GA-356A DISCH. |  |  |  |  | D-1-IPA-3004 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 3.44 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.7 | 0 | 120 | 0 |  | O
40 | PG-1356B | GA-356B DISCH. |  |  |  |  | D-1-IPA-3004 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 3.44 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.7 | 0 | 120 | 0 |  | O
41 | PG-1357 | DA-350 EQUIP. |  |  |  |  | D-1-IPA-3004 |  |  | PDI / PIP |  | M | -1 ~ 2 |  |  |  | 0.4 |  |  | 90 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 120 | 0 |  | O
42 | PG-1358 | N2 TO FLARE |  |  |  |  | D-1-IPA-3005 |  |  | NL |  | V | -1 ~ 1 |  |  |  | 0.01 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
43 | PG-1361 | LP TO EA-361 |  |  |  |  | D-1-IPA-3006 |  |  | SL |  | S | 0 ~ 5 |  |  |  | 2.5 |  |  | 138.2 |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 6 | 0 | 200 | 0 |  | O
44 | PG-1362 | EA-361 EQUIP. |  |  |  |  | D-1-IPA-3006 |  |  | SL |  | S | 0 ~ 5 |  |  |  | 2.5 |  |  | 138.2 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 6 | 0 | 200 | 0 |  | O
45 | PG-1363A | GA-363A DISCH. |  |  |  |  | D-1-IPA-3006 |  |  | PPW |  | L | 0 ~ 10 |  |  |  | 6.49 |  |  | 108 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.4 | 0 | 140 | 0 |  | O
46 | PG-1363B | GA-363B DISCH. |  |  |  |  | D-1-IPA-3006 |  |  | PPW |  | L | 0 ~ 10 |  |  |  | 6.49 |  |  | 108 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.4 | 0 | 140 | 0 |  | O
47 | PG-1364 | DA-360 EQUIP. |  |  |  |  | D-1-IPA-3006 |  |  | PPW |  | L | -1 ~ 1 |  |  |  | 0.3 |  |  | 108 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 140 | 0 |  | O
48 | PG-1365A | GA-365A DISCH. |  |  |  |  | D-1-IPA-3006 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 4.76 |  |  | 74 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.9 | 0 | 120 | 0 |  | O
49 | PG-1365B | GA-365B DISCH. |  |  |  |  | D-1-IPA-3006 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 4.76 |  |  | 74 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.9 | 0 | 120 | 0 |  | O
50 | PG-1367 | EA-365 EQUIP. |  |  |  |  | D-1-IPA-3006 |  |  | PDI |  | V | -1 ~ 1 |  |  |  | 0.05 |  |  | 80.3 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 6 | 0 | 120 | 0 |  | O
51 | PG-1371 | FA-371 EQUIP. |  |  |  |  | D-1-IPA-3007 |  |  | VA |  | V | -1 ~ 1 |  |  |  | 0.03 / -0.02 |  |  | 50 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 |  | O
52 | PG-1371A | GA-371A DISCH. |  |  |  |  | D-1-IPA-3007 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.19 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.1 | 0 | 120 | 0 |  | O
53 | PG-1371B | GA-371B DISCH. |  |  |  |  | D-1-IPA-3007 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.19 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.1 | 0 | 120 | 0 |  | O
54 | PG-1372A | FA-373A EQUIP. |  |  |  |  | D-1-IPA-3007 |  |  | GFL |  | V | -1 ~ 1 |  |  |  | 0.03 / -0.02 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 |  | O
55 | PG-1372B | FA-373B EQUIP. |  |  |  |  | D-1-IPA-3007 |  |  | GFL |  | V | -1 ~ 1 |  |  |  | 0.03 / -0.02 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 |  | O
56 | PG-1372C | FA-373C EQUIP. |  |  |  |  | D-1-IPA-3007 |  |  | GFL |  | V | -1 ~ 1 |  |  |  | 0.03 / -0.02 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 |  | O
57 | PG-1373A | GA-373A DISCH. |  |  |  |  | D-1-IPA-3007 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.66 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.6 | 0 | 120 | 0 |  | O
58 | PG-1373B | GA-373B DISCH. |  |  |  |  | D-1-IPA-3007 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.66 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.6 | 0 | 120 | 0 |  | O
59 | PG-1374 | N2 TO FLARE |  |  |  |  | D-1-IPA-3007 |  |  | NL |  | V | -1 ~ 1 |  |  |  | 0.01 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
60 | PG-1374A | GA-374A DISCH. |  |  |  |  | D-1-IPA-3007 |  |  | PDI |  | L | 0 ~ 5 |  |  |  | 2.1 |  |  | 70 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 3.5 | 0 | 150 | 0 |  | O
61 | PG-1374B | GA-374B DISCH. |  |  |  |  | D-1-IPA-3007 |  |  | PDI |  | L | 0 ~ 5 |  |  |  | 2.1 |  |  | 70 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 3.5 | 0 | 150 | 0 |  | O
62 | PG-1375 | FD-373A/B DISCH. |  |  |  |  | D-1-IPA-3007 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.66 |  |  | 40 |  |  |  | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.6 | 0 | 100 | 0 |  | O
63 | PG-1376 | N2 TO FLARE |  |  |  |  | D-1-IPA-3007 |  |  | NL |  | V | 0 ~ 5 |  |  |  | 2 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
64 | PG-1381 | N2 TO FA-371 |  |  |  |  | D-1-IPA-3007 |  |  | NL |  | V | 0 ~ 10 |  |  |  | 2 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
65 | PG-1401 | IPA TO LOADING SYSTEM |  |  |  |  | D-1-IPA-4004 |  |  | PIP |  | L | 0 ~ 3 |  |  |  | 1.54 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 3.41 | 0 | 120 | 0 |  | O
66 | PG-1402 | N2 TO LOADING SYSTEM |  |  |  |  | D-1-IPA-4004 |  |  | NL |  | V | 0 ~ 15 |  |  |  | 7 |  |  | AMB |  |  | D | LF | CONN. : 1/2" 150# RF | 1" 150# RF | BB55C1 | 150# RF | 10 | 0 | 60 | 0 |  | O
67 | PG-1403 | IPA TO LOADING SYSTEM |  |  |  |  | D-1-IPA-4004 |  |  | PIP |  | L | 0 ~ 3 |  |  |  | 1.54 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 3.41 | 0 | 120 | 0 |  | O
68 | PG-1404 | N2 TO LOADING SYSTEM |  |  |  |  | D-1-IPA-4004 |  |  | NL |  | V | 0 ~ 15 |  |  |  | 7 |  |  | AMB |  |  | D | LF | CONN. : 1/2" 150# RF | 1" 150# RF | BB55C1 | 150# RF | 10 | 0 | 60 | 0 |  | O
69 | PG-1405 | IPE TO LOADING SYSTEM |  |  |  |  | D-1-IPA-4004 |  |  | PDI |  | L | 0 ~ 5 |  |  |  | 2.83 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.24 | 0 | 120 | 0 |  | O
70 | PG-1406 | N2 TO LOADING SYSTEM |  |  |  |  | D-1-IPA-4004 |  |  | NL |  | V | 0 ~ 15 |  |  |  | 7 |  |  | AMB |  |  | D | LF | CONN. : 1/2" 150# RF | 1" 150# RF | BB55C1 | 150# RF | 10 | 0 | 60 | 0 |  | O
71 | PG-1420A | GA-420A DISCH. |  |  |  |  | D-1-IPA-4002 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.25 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.52 | 0 | 120 | 0 |  | O
72 | PG-1420B | GA-420B DISCH. |  |  |  |  | D-1-IPA-4002 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.25 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.52 | 0 | 120 | 0 |  | O
73 | PG-1422 | N2 TO FB-420A/B |  |  |  |  | D-1-IPA-4002 |  |  | NL |  | V | 0 ~ 5 |  |  |  | 2 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
74 | PG-1424 | FD-420A INLET |  |  |  |  | D-1-IPA-4002 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.25 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.52 | 0 | 120 | 0 |  | O
75 | PG-1425 | FD-420A OUTLET |  |  |  |  | D-1-IPA-4002 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.1 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.5 | 0 | 120 | 0 |  | O
76 | PG-1426 | FD-420B INLET |  |  |  |  | D-1-IPA-4002 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.25 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.52 | 0 | 120 | 0 |  | O
77 | PG-1427 | FD-420B OUTLET |  |  |  |  | D-1-IPA-4002 |  |  | PIP |  | L | 0 ~ 10 |  |  |  | 4.1 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.5 | 0 | 120 | 0 |  | O
78 | PG-1430 | GA-430 DISCH. |  |  |  |  | D-1-IPA-4003 |  |  | PPA |  | L | 0 ~ 3 |  |  |  | 1.2 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | FF55C | 150# RF | 2.1 | 0 | 120 | 0 |  | O
79 | PG-1431 | GA-431 DISCH. |  |  |  |  | D-1-IPA-4003 |  |  | PS |  | L | 0 ~ 5 |  |  |  | 2.19 |  |  | 20 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.9 | 0 | 120 | 0 |  | O
80 | PG-1435 | NS FROM GA-982XC |  |  |  |  | D-1-IPA-4003 |  |  | NS |  | L | 0 ~ 15 |  |  |  | 7.7 |  |  | 40 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10 | 0 | 120 | 0 |  | O
81 | PG-1781 | FA-781A EQUIP. |  |  |  |  | D-1-IPA-7001 |  |  | PDI |  | V | -1 ~ 1 |  |  |  | ATM |  |  | AMB |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 |  | O
82 | PG-1782 | FA-781B EQUIP. |  |  |  |  | D-1-IPA-7001 |  |  | PDI |  | V | -1 ~ 1 |  |  |  | ATM |  |  | AMB |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 |  | O
83 | PG-1783 | GA-781A DISCH. |  |  |  |  | D-1-IPA-7001 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 4.57 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.58 | 0 | 120 | 0 |  | O
84 | PG-1784 | GA-781B DISCH. |  |  |  |  | D-1-IPA-7001 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 4.57 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.58 | 0 | 120 | 0 |  | O
85 | PG-1808 | SL TO CONSUMERS |  |  |  |  | D-1-IPA-8011 |  |  | SL |  | S | 0 ~ 10 |  |  |  | 3.5 |  |  | 148 |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 6 | F.V | 200 | 0 |  | O
86 | PG-1815A | GA-815A DISCH. |  |  |  |  | D-1-IPA-8011 |  |  | CL |  | L | 0 ~ 20 |  |  |  | 11.3 |  |  | 100 |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 14.3 | 0 | 150 | 0 |  | O
87 | PG-1815B | GA-815B DISCH. |  |  |  |  | D-1-IPA-8011 |  |  | CL |  | L | 0 ~ 20 |  |  |  | 11.3 |  |  | 100 |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 14.3 | 0 | 150 | 0 |  | O
88 | PG-1820A | GA-820A DISCH. |  |  |  |  | D-1-IPA-8012 |  |  | WCS |  | L | 0 ~ 10 |  |  |  | 6.5 |  |  | 32 |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 8.5 | 0 | AMB | 0 |  | O
89 | PG-1820B | GA-820B DISCH. |  |  |  |  | D-1-IPA-8012 |  |  | WCS |  | L | 0 ~ 10 |  |  |  | 6.5 |  |  | 32 |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 8.5 | 0 | AMB | 0 |  | O
90 | PG-1821 | WCR TO EF-820 |  |  |  |  | D-1-IPA-8012 |  |  | WCR |  | L | 0 ~ 5 |  |  |  | 2 |  |  | 42 |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 7 | 0 | 50 | 0 |  | O
91 | PG-1830A | GA-830A DISCH. |  |  |  |  | D-1-IPA-8013 |  |  | WT |  | L | 0 ~ 10 |  |  |  | 6.4 |  |  | 120 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 9.5 | 0 | 180 | 0 |  | O
92 | PG-1830B | GA-830B DISCH. |  |  |  |  | D-1-IPA-8013 |  |  | WT |  | L | 0 ~ 10 |  |  |  | 6.4 |  |  | 120 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 9.5 | 0 | 180 | 0 |  | O
93 | PG-1831 | FA-830X EQUIP. |  |  |  |  | D-1-IPA-8013 |  |  | WT |  | L | -1 ~ 1 |  |  |  | 0.25 |  | 120 | 105 |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 4 | F.V | 180 | 0 |  | O
94 | PG-1835A | GA-835A DISCH. |  |  |  |  | D-1-IPA-8013 |  |  | WT |  | L | 0 ~ 10 |  |  |  | 6.7 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.7 | 0 | 150 | 0 |  | O
95 | PG-1835B | GA-835B DISCH. |  |  |  |  | D-1-IPA-8013 |  |  | WT |  | L | 0 ~ 10 |  |  |  | 6.7 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.7 | 0 | 150 | 0 |  | O
96 | PG-1838 | DA-830 EQUIP. |  |  |  |  | D-1-IPA-8013 |  |  | WT |  | L | -1 ~ 1 |  |  | 1 | 0.25 |  |  | 105 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 4 | F.V | 180 | 0 |  | O
97 | PG-1839 | WT FROM BL |  |  |  |  | D-1-IPA-8013 |  |  | WT |  | L | 0 ~ 10 |  |  |  | 5 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7 | 0 | AMB | 0 |  | O
98 | PG-1842 | AI FROM BL |  |  |  |  | D-1-IPA-8014 |  |  | AI |  | V | 0 ~ 15 |  |  |  | 7 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BC55C | 150# RF | 10 | 0 | 60 | 0 |  | O
99 | PG-1843 | FA-841 TO AI HEADER |  |  |  |  | D-1-IPA-8014 |  |  | AI |  | V | 0 ~ 15 |  |  |  | 7 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BC55C | 150# RF | 10 | 0 | 60 | 0 |  | O
100 | PG-1844 | FA-841 EQUIP. |  |  |  |  | D-1-IPA-8014 |  |  | AI |  | V | 0 ~ 15 |  |  |  | 7 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 10 | 0 | 60 | 0 |  | O
101 | PG-1847 | NL TO NL HEADER |  |  |  |  | D-1-IPA-8014 |  |  | NL |  | V | 0 ~ 15 |  |  |  | 7 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
102 | PG-1849 | NL FROM BL |  |  |  |  | D-1-IPA-8014 |  |  | NL |  | V | 0 ~ 15 |  |  |  | 7 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
103 | PG-1850A | GA-840A DISCH. |  |  |  |  | D-1-IPA-8024 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.5 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 7.9 | 0 | 70 | 0 |  | O
104 | PG-1850B | GA-840B DISCH. |  |  |  |  | D-1-IPA-8024 |  |  | PDI |  | L | 0 ~ 10 |  |  |  | 5.5 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 7.9 | 0 | 70 | 0 |  | O
105 | PG-1851A | GA-850A DISCH. |  |  |  |  | D-1-IPA-8024 |  |  | WW |  | L | 0 ~ 3 |  |  |  | 1.55 |  |  | 30 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 2.7 | 0 | 70 | 0 |  | O
106 | PG-1851B | GA-850B DISCH. |  |  |  |  | D-1-IPA-8024 |  |  | WW |  | L | 0 ~ 3 |  |  |  | 1.55 |  |  | 30 |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 2.7 | 0 | 70 | 0 |  | O
107 | PG-1860 | FA-860 EQUIP. |  |  |  |  | D-1-IPA-8023-1 |  |  | VFL |  | V | -1 ~ 1 |  |  |  | 0.3 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 300 | 0 |  | O
108 | PG-1860A | GA-860A DISCH. |  |  |  |  | D-1-IPA-8023-1 |  |  | PDI |  | L | 0 ~ 5 |  |  |  | 2.18 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 4 | 0 | 300 | 0 |  | O
109 | PG-1860B | GA-860B DISCH. |  |  |  |  | D-1-IPA-8023-1 |  |  | PDI |  | L | 0 ~ 5 |  |  |  | 2.18 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 4 | 0 | 300 | 0 |  | O
110 | PG-1870 | AP FROM BL |  |  |  |  | D-1-IPA-8015 |  |  | AP |  | V | 0 ~ 15 |  |  |  | 7 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 |  | O
111 | PG-1880 | WF FROM BL |  |  |  |  | D-1-IPA-8022 |  |  | WR |  | L | 0 ~ 10 |  |  |  | 4 |  |  | AMB |  |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 7 | 0 | 60 | 0 |  | O
112 | PG-1900A | GA-900A DISCH. |  |  |  |  | D-1-IPA-8024 |  |  | WW |  | L | HOLD |  |  |  | HOLD |  |  | HOLD |  |  | D | LF | HOLD | 1" 150# RF | -- | 150# RF | -- | -- | -- | -- |  | X
113 | PG-1900B | GA-900B DISCH. |  |  |  |  | D-1-IPA-8024 |  |  | WW |  | L | HOLD |  |  |  | HOLD |  |  | HOLD |  |  | D | LF | HOLD | 1" 150# RF | -- | 150# RF | -- | -- | -- | -- |  | X
#N/A

## Flanged_PG_300
[9 empty] | DATA SHEET | [13 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  |  | : M0556 |  |  |  |  | LOCAL INSTRUMENT (PRESSURE GAUGE) | [13 empty] | DS-IN-F-018A
| PROJECT NAME |  |  | : IPA 50KTA Exp. | [18 empty] | REV. NO.
| LOCATION |  |  | : Yeosu, Korea | [18 empty] | A
GENERAL | [12 empty] | ELEMENT
1. | TYPE |  | ■ | DIRECT READING |  |  |  | □ | RECEIVER |  |  |  | 13. | ELEMENT TYPE |  |  |  | ■ | BOURDON |  |  |  | □ | BELLOWS
2. | MOUNTING |  | ■ | LOCAL |  | □ | SURFACE |  |  | □ | FLUSH |  | 14. | ELEMENT MAT'L |  |  |  | ■ | 316SS |  | □ | OTHER
3. | DIAL SIZE |  | ■ | 100 mm |  | □ | 150 mm |  |  | □ | 4 1/2" |  | 15. | SOCKET MAT'L |  |  |  | ■ | 316SS |  | □ | OTHER
4. | DIAL COLOR |  | ■ | WHITE BACK GROUND BLACK LETTERS | [8 empty] | 16. | CONNECTION |  |  |  | ■ | 1/2" NPT(M) |  | □ | OTHER
5. | CASE MAT'L |  | ■ | 304SS |  | □ | 316SS |  |  | □ | ALUMINUM |  | 17. | MOVEMENT |  |  |  | ■ | 304SS |  | □ | OTHER
6. | RING TYPE |  | □ | SCREWED |  | ■ | BAYONET |  |  | □ | SLIP |  | 18. | DIAPHRAGM SEAL
7. | ACCURACY |  | ■ | ± 1.0 % OF F.S |  |  | □ | OTHER | [6 empty] | DIAPHRAGM MAT'L |  |  |  |  | 316LSS
8. | OVERRANGE PROTECTION |  |  |  | ■ | 1.3 x F.S | [8 empty] | UPPER FLANGE MAT'L |  |  |  |  | 316SS
9. | BLOW OUT PROTECTION |  |  |  | ■ | BACK |  |  |  | □ | SIDE |  |  |  | UNDER FLANGE MAT'L |  |  |  |  | 316SS
[5 empty] | ■ | SOLID FRONT ( ≥ 30 kg/cm2g) | [8 empty] | FILL FLUID |  |  |  |  | SILICONE OIL
10. | LENS |  | ■ | SAFETY GLASS |  |  | □ | OTHER | [6 empty] | PROCESS CONNECTION |  |  |  |  | 1" 300# RF
OPTION | [12 empty] | 19. | MFR & MODEL NO.
11. | ACCESSORIES |  | □ | SYPHONE (S) |  |  |  | □ | SNUBBER (B)
| (SEE LIST) |  | ■ | DAMPENER (D) |  |  |  | □ | SOLID FRONT (SF) |  |  |  | NOTE
|  |  | ■ | LIQUID FILLED (LF) | [8 empty] | 1. | Syphone Tube shall be min. 80 sch.
12. | MANDATORY OPTION | [11 empty] | 2. | Vendor shall be take consideration the service condition for seal fluid
|  |  | □ | GAUGE UNION (G) |  |  |  | □ | MANIFOLD VALVE (M) |  |  |  |  | and accessories.
|  |  | ■ | COOLING FIN (C) - MFR. STD.
NO | TAG NO | SERVICE |  |  |  |  | P&ID NO |  |  | FLUID |  | STATE | RANGE (kg/cm2g) |  |  | PRESSURE (kg/cm2g) |  |  | TEMP. (deg C) |  |  | ACCESSORIES |  |  | REMARK
[29 empty] | DESIGN PRESS. |  | DESIGN TEMP. |  | REMARK
[16 empty] | MAX | NOR |  | MAX | NOR | [5 empty] | TYPE | SPEC. | RATING | INT. | EXT. | HOT | COLD
1 | PG-1215A | GA-215A DISCH. |  |  |  |  | D-1-IPA-2001 |  |  | PPR |  | L | 0 ~ 40 |  |  |  | 22.9 |  |  | 40 |  |  | D | LF |  | 1" 300# RF | BB57C1 | 300# RF | 31.7 | 0 | 120 | 0 |  | O
2 | PG-1215B | GA-215B DISCH. |  |  |  |  | D-1-IPA-2001 |  |  | PPR |  | L | 0 ~ 40 |  |  |  | 22.9 |  |  | 40 |  |  | D | LF |  | 1" 300# RF | BB57C1 | 300# RF | 31.7 | 0 | 120 | 0 |  | O
3 | PG-1220 | DA-220 EQUIP. |  |  |  |  | D-1-IPA-2002 |  |  | PHE / PPR |  | M | 0 ~ 30 |  |  |  | 19.6 |  |  | 52 |  |  |  | LF |  | 1" 150# RF | - | 300# RF | 25 | 0 | 120 | 0 |  | O
4 | PG-1227 | GA-226A/B SUCTION |  |  |  |  | D-1-IPA-2002 |  |  | PPR |  | L | 0 ~ 30 |  |  |  | 19.32 |  |  | 47.2 |  |  |  | LF |  | 1" 300# RF | BB57C1 | 300# RF | 25 | 0 | 120 | -46 |  | O
5 | PG-1311 | MP STEAM TO EA-311 |  |  |  |  | D-1-IPA-3001 |  |  | SM |  | S | 0 ~ 15 |  |  |  | 8 |  |  | 174.5 |  |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | 0 | 250 | 0 |  | O
6 | PG-1321 | MP STEAM TO EA-321 |  |  |  |  | D-1-IPA-3002 |  |  | SM |  | S | 0 ~ 15 |  |  |  | 8 |  |  | 174.5 |  |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | 0 | 250 | 0 |  | O
7 | PG-1352 | EA-352 EQUIP. |  |  |  |  | D-1-IPA-3004 |  |  | SM |  | S | 0 ~ 15 |  |  |  | 8 |  |  | 174.5 |  |  |  | LF |  | 1" 300# RF | - | 300# RF | 12 | 0 | 250 | 0 |  | O
8 | PG-1354 | EA-352 EQUIP. |  |  |  |  | D-1-IPA-3004 |  |  | SM |  | S | 0 ~ 15 |  |  |  | 8 |  |  | 174.5 |  |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | 0 | 250 | 0 |  | O
9 | PG-1401A | GA-410A DISCH. |  |  |  |  | B-1-9112 |  |  | PPR |  | L | 0 ~ 30 |  |  |  | 18 |  | 35 | 25 |  |  | D | LF |  | 1" 300# RF | BA8A | 300# RF | 25 | 0 | 65 | -15 |  | O
10 | PG-1401B | GA-410B DISCH. |  |  |  |  | B-1-9112 |  |  | PPR |  | L | 0 ~ 30 |  |  |  | 18 |  | 35 | 25 |  |  | D | LF |  | 1" 300# RF | BA8A | 300# RF | 25 | 0 | 65 | -15 |  | O
11 | PG-1412 | PPR FROM GA-410A/B |  |  |  |  | D-1-IPA-4001 |  |  | PPR |  | L | 0 ~ 30 |  |  |  | 18.01 |  | 35 | 25 |  |  |  | LF |  | 1" 300# RF | BA8A | 300# RF | 25 | 0 | 65 | -15 |  | O
12 | PG-1415 | FA-415 EQUIP. |  |  |  |  | D-1-IPA-4001 |  |  | VFL |  | V | 0 ~ 30 |  |  |  | 15.2 |  |  | AMB |  |  |  | LF |  | 1" 300# RF | - | 300# RF | 25 | 0 | 65 | -15 |  | O
13 | PG-1418A | GA-415A SUCTION |  |  |  |  | D-1-IPA-4001 |  |  | PPF |  | L | 0 ~ 30 |  |  |  | 15.4 |  | 35 | 25 |  |  |  | LF |  | 1" 300# RF | BB57C1 | 300# RF | 25 | 0 | 65 | -15 |  | O
14 | PG-1418B | GA-415B SUCTION |  |  |  |  | D-1-IPA-4001 |  |  | PPF |  | L | 0 ~ 30 |  |  |  | 15.4 |  | 35 | 25 |  |  |  | LF |  | 1" 300# RF | BB57C1 | 300# RF | 25 | 0 | 65 | -15 |  | O
15 | PG-1807 | SM TO CONSUMERS |  |  |  |  | D-1-IPA-8011 |  |  | SM |  | S | 0 ~ 15 |  |  |  | 9.8 |  |  | 198 |  |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | F.V | 250 | 0 |  | O
16 | PG-1833 | DA-830 TO EA-830X INLET |  |  |  |  | D-1-IPA-8013 |  |  | CM |  | L | 0 ~ 10 |  |  | 8 | 2.5 |  | 193.5 | 143.5 |  |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | F.V | 250 | 0 |  | O
17 | PG-1834 | EA-830X OUTLET TO CM |  |  |  |  | D-1-IPA-8013 |  |  | CM |  | L | 0 ~ 10 |  |  | 8 | 2.5 |  | 193.5 | 143.5 |  |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | F.V | 250 | 0 |  | O
18 | PG-1841 | FA-845A/B TO FA-841 |  |  |  |  | D-1-IPA-8014 |  |  | AI |  | V | 0 ~ 15 |  |  |  | 10 |  |  | AMB |  |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 13 | 0 | 65 | 0 |  | O

## Flanged_PG_600
[9 empty] | DATA SHEET | [13 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  |  | : M0556 |  |  |  |  | LOCAL INSTRUMENT (PRESSURE GAUGE) | [13 empty] | DS-IN-F-018A
| PROJECT NAME |  |  | : IPA 50KTA Exp. | [18 empty] | REV. NO.
| LOCATION |  |  | : Yeosu, Korea | [18 empty] | A
GENERAL | [12 empty] | ELEMENT
1. | TYPE |  | ■ | DIRECT READING |  |  |  | □ | RECEIVER |  |  |  | 13. | ELEMENT TYPE |  |  |  | ■ | BOURDON |  |  |  | □ | BELLOWS
2. | MOUNTING |  | ■ | LOCAL |  | □ | SURFACE |  |  | □ | FLUSH |  | 14. | ELEMENT MAT'L |  |  |  | ■ | 316SS |  | □ | OTHER
3. | DIAL SIZE |  | ■ | 100 mm |  | □ | 150 mm |  |  | □ | 4 1/2" |  | 15. | SOCKET MAT'L |  |  |  | ■ | 316SS |  | □ | OTHER
4. | DIAL COLOR |  | ■ | WHITE BACK GROUND BLACK LETTERS | [8 empty] | 16. | CONNECTION |  |  |  | ■ | 1/2" NPT(M) |  | □ | OTHER
5. | CASE MAT'L |  | ■ | 304SS |  | □ | 316SS |  |  | □ | ALUMINUM |  | 17. | MOVEMENT |  |  |  | ■ | 304SS |  | □ | OTHER
6. | RING TYPE |  | □ | SCREWED |  | ■ | BAYONET |  |  | □ | SLIP |  | 18. | DIAPHRAGM SEAL
7. | ACCURACY |  | ■ | ± 1.0 % OF F.S |  |  | □ | OTHER | [6 empty] | DIAPHRAGM MAT'L |  |  |  |  | 316LSS
8. | OVERRANGE PROTECTION |  |  |  | ■ | 1.3 x F.S | [8 empty] | UPPER FLANGE MAT'L |  |  |  |  | 316SS
9. | BLOW OUT PROTECTION |  |  |  | ■ | BACK |  |  |  | □ | SIDE |  |  |  | UNDER FLANGE MAT'L |  |  |  |  | 316SS
[5 empty] | ■ | SOLID FRONT ( ≥ 30 kg/cm2g) | [8 empty] | FILL FLUID |  |  |  |  | SILICONE OIL
10. | LENS |  | ■ | SAFETY GLASS |  |  | □ | OTHER | [6 empty] | PROCESS CONNECTION |  |  |  |  | 1" 600# RF
OPTION | [12 empty] | 19. | MFR & MODEL NO.
11. | ACCESSORIES |  | □ | SYPHONE (S) |  |  |  | □ | SNUBBER (B)
| (SEE LIST) |  | ■ | DAMPENER (D) |  |  |  | ■ | SOLID FRONT (SF) |  |  |  | NOTE
|  |  | ■ | LIQUID FILLED (LF) | [8 empty] | 1. | Syphone Tube shall be min. 80 sch.
12. | MANDATORY OPTION | [11 empty] | 2. | Vendor shall be take consideration the service condition for seal fluid
|  |  | □ | GAUGE UNION (G) |  |  |  | □ | MANIFOLD VALVE (M) |  |  |  |  | and accessories.
|  |  | ■ | COOLING FIN (C) - MFR. STD.
NO | TAG NO | SERVICE |  |  |  |  | P&ID NO |  |  | FLUID |  | STATE | RANGE (kg/cm2g) |  |  | PRESSURE (kg/cm2g) |  |  | TEMP. (deg C) |  |  | ACCESSORIES |  |  |  | REMARK
[30 empty] | DESIGN PRESS. |  | DESIGN TEMP. |  | REMARK
[16 empty] | MAX | NOR |  | MAX | NOR | [6 empty] | TYPE | SPEC. | RATING | INT. | EXT. | HOT | COLD
1 | PG-1110A | GA-110A TO DC-110 |  |  |  |  | D-1-IPA-1001 |  |  | PPA |  | L | 0 ~ 80 |  |  |  | 50.34 |  |  | 20 |  |  | D | LF | SF |  | 1" 600# RF | FF59C | 600# RF | 60 | 0 | 120 | 0 |  | O
2 | PG-1110B | GA-110A TO DC-110 |  |  |  |  | D-1-IPA-1001 |  |  | PPA |  | L | 0 ~ 80 |  |  |  | 50.34 |  |  | 20 |  |  | D | LF | SF |  | 1" 600# RF | FF59C | 600# RF | 60 | 0 | 120 | 0 |  | O
3 | PG-1112A | GA-112A DISCH. |  |  |  |  | D-1-IPA-1001 |  |  | PCS |  | L | 0 ~ 80 |  |  |  | 51.2 |  |  | 20 |  |  | D | LF | SF |  | 1" 150# RF | BB59C1 | 600# RF | 60 | 0 | 120 | 0 |  | O
4 | PG-1112B | GA-112B DISCH. |  |  |  |  | D-1-IPA-1001 |  |  | PCS |  | L | 0 ~ 80 |  |  |  | 51.2 |  |  | 20 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 60 | 0 | 120 | 0 |  | O
5 | PG-1113 | GA-113 DISCH. |  |  |  |  | D-1-IPA-1001 |  |  | PCS |  | L | 0 ~ 80 |  |  |  | 51.3 |  |  | 20 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 60 | 0 | 120 | 0 |  | O
6 | PG-1115A | GA-115A DISCH. |  |  |  |  | D-1-IPA-1001 |  |  | PPW |  | L | 0 ~ 80 |  |  |  | 49.8 |  |  | 103 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 61.1 | 0 | 160 | 0 |  | O
7 | PG-1115B | GA-115B DISCH. |  |  |  |  | D-1-IPA-1001 |  |  | PPW |  | L | 0 ~ 80 |  |  |  | 49.8 |  |  | 103 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 61.1 | 0 | 160 | 0 |  | O
8 | PG-1140 | GB-120 DISCH. |  |  |  |  | D-1-IPA-1002 |  |  | PRG |  | V | 0 ~ 80 |  |  |  | 47 |  |  | 125 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 55 | 0 | 150 | 0 |  | O
9 | PG-1141 | FA-121 EQUIP. |  |  |  |  | D-1-IPA-1002 |  |  | PRG |  | V | 0 ~ 70 |  |  |  | 44.8 |  |  | 120 |  |  |  | LF | SF |  | 1" 600# RF | - | 600# RF | 55 | 0 | 180 | 0 |  | O
10 | PG-1142 | GB-120 SUCTION |  |  |  |  | D-1-IPA-1002 |  |  | PRG |  | V | 0 ~ 50 |  |  |  | 31 |  |  | 94 |  |  |  | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 55 | 0 | 150 | 0 |  | O
11 | PG-1145A | GA-145A DISCH. |  |  |  |  | D-1-IPA-1002 |  |  | PWW |  | L | 0 ~ 70 |  |  |  | 39.25 |  |  | 95 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 48.1 | 0 | 150 | 0 |  | O
12 | PG-1145B | GA-145B DISCH. |  |  |  |  | D-1-IPA-1002 |  |  | PWW |  | L | 0 ~ 70 |  |  |  | 39.25 |  |  | 95 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 48.1 | 0 | 150 | 0 |  | O
13 | PG-1146 | FC-120 EQUIP. |  |  |  |  | D-1-IPA-1002 |  |  | PRG |  | V | 0 ~ 50 |  |  |  | 31 |  |  | 94 |  |  | D | LF |  |  | 1" 600# RF | BB59C1 | 600# RF | 55 | 0 | 150 | 0
14 | PG-1225A | GA-225A DISCH. |  |  |  |  | D-1-IPA-2002 |  |  | PPR |  | L | 0 ~ 40 |  |  |  | 22.15 |  |  | 45 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 28.3 | 0 | 120 | -46 |  | O
15 | PG-1225B | GA-225B DISCH. |  |  |  |  | D-1-IPA-2002 |  |  | PPR |  | L | 0 ~ 40 |  |  |  | 22.15 |  |  | 45 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 28.3 | 0 | 120 | -46 |  | O
16 | PG-1226A | GA-226A DISCH. |  |  |  |  | D-1-IPA-2002 |  |  | PPR |  | L | 0 ~ 80 |  |  |  | 50.9 |  |  | 47.2 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 58.9 | 0 | 120 | -46 |  | O
17 | PG-1226B | GA-226B DISCH. |  |  |  |  | D-1-IPA-2002 |  |  | PPR |  | L | 0 ~ 80 |  |  |  | 50.9 |  |  | 47.2 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 58.9 | 0 | 120 | -46 |  | O
18 | PG-1316A | GA-316A DISCH. |  |  |  |  | D-1-IPA-3001 |  |  | PE |  | L | 0 ~ 80 |  |  |  | 50.83 |  |  | 45 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 63.5 | 0 | 120 | 0 |  | O
19 | PG-1316B | GA-316B DISCH. |  |  |  |  | D-1-IPA-3001 |  |  | PE |  | L | 0 ~ 80 |  |  |  | 50.83 |  |  | 45 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 63.5 | 0 | 120 | 0 |  | O
20 | PG-1416A | GA-415A DISCH. |  |  |  |  | D-1-IPA-4001 |  |  | PPF |  | L | 0 ~ 80 |  |  |  | 49.7 |  | 35 | 25 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 55 | 0 | 65 | -15 |  | O
21 | PG-1416B | GA-415B DISCH. |  |  |  |  | D-1-IPA-4001 |  |  | PPF |  | L | 0 ~ 80 |  |  |  | 49.7 |  | 35 | 25 |  |  | D | LF | SF |  | 1" 600# RF | BB59C1 | 600# RF | 55 | 0 | 65 | -15 |  | O
22 | PG-1802 | BFW FROM BL |  |  |  |  | D-1-IPA-8011 |  |  | BFW |  | L | 0 ~ 100 |  |  |  | 64.6 |  |  | 80 |  |  | D | LF |  |  | 1" 600# RF | BA59C | 600# RF | 66 | 0 | 300 | 0
23 | PG-1845 | FA-845A EQUIP. |  |  |  |  | D-1-IPA-8014 |  |  | NH |  | V | 0 ~ 110 |  |  |  | 70 |  |  | AMB |  |  |  | LF | SF |  | 1" 600# RF | - | 600# RF | 77 | 0 | 65 | 0 |  | O

## NPT_PG
[9 empty] | DATA SHEET | [13 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  |  | : M0556 |  |  |  |  | LOCAL INSTRUMENT (PRESSURE GAUGE) | [13 empty] | DS-IN-F-018A
| PROJECT NAME |  |  | : IPA 50KTA Exp. | [18 empty] | REV. NO.
| LOCATION |  |  | : Yeosu, Korea | [18 empty] | A
GENERAL | [12 empty] | ELEMENT
1. | TYPE |  | ■ | DIRECT READING |  |  |  | □ | RECEIVER |  |  |  | 13. | ELEMENT TYPE |  |  |  | ■ | BOURDON |  |  |  | □ | BELLOWS
2. | MOUNTING |  | ■ | LOCAL |  | □ | SURFACE |  |  | □ | FLUSH |  | 14. | ELEMENT MAT'L |  |  |  | ■ | 316SS |  | □ | OTHER
3. | DIAL SIZE |  | ■ | 100 mm |  | □ | 150 mm |  |  | □ | 4 1/2" |  | 15. | SOCKET MAT'L |  |  |  | ■ | 316SS |  | □ | OTHER
4. | DIAL COLOR |  | ■ | WHITE BACK GROUND BLACK LETTERS | [8 empty] | 16. | CONNECTION |  |  |  | ■ | 1/2" NPT(M) |  | □ | OTHER
5. | CASE MAT'L |  | ■ | 304SS |  | □ | 316SS |  |  | □ | ALUMINUM |  | 17. | MOVEMENT |  |  |  | ■ | 304SS |  | □ | OTHER
6. | RING TYPE |  | □ | SCREWED |  | ■ | BAYONET |  |  | □ | SLIP |  | 18. | DIAPHRAGM SEAL
7. | ACCURACY |  | ■ | ± 1.0 % OF F.S |  |  | □ | OTHER | [6 empty] | DIAPHRAGM MAT'L |  |  |  |  | -
8. | OVERRANGE PROTECTION |  |  |  | ■ | 1.3 x F.S | [8 empty] | UPPER FLANGE MAT'L |  |  |  |  | -
9. | BLOW OUT PROTECTION |  |  |  | ■ | BACK |  |  |  | □ | SIDE |  |  |  | UNDER FLANGE MAT'L |  |  |  |  | -
[5 empty] | ■ | SOLID FRONT ( ≥ 30 kg/cm2g) | [8 empty] | FILL FLUID |  |  |  |  | -
10. | LENS |  | ■ | SAFETY GLASS |  |  | □ | OTHER | [6 empty] | PROCESS CONNECTION |  |  |  |  | 1/2" NPT
OPTION | [12 empty] | 19. | MFR & MODEL NO.
11. | ACCESSORIES |  | ■ | SYPHONE (S) |  |  |  | □ | SNUBBER (B)
| (SEE LIST) |  | □ | DAMPENER (D) |  |  |  | ■ | SOLID FRONT (SF) |  |  |  | NOTE
|  |  | ■ | LIQUID FILLED (LF) | [8 empty] | 1. | Syphone Tube shall be min. 80 sch.
12. | MANDATORY OPTION | [11 empty] | 2. | Vendor shall be take consideration the service condition for seal fluid
|  |  | □ | GAUGE UNION (G) |  |  |  | ■ | MANIFOLD VALVE (M) |  |  |  |  | and accessories.
|  |  | ■ | COOLING FIN (C) - MFR. STD.
NO | TAG NO | SERVICE |  |  |  |  | P&ID NO |  |  | FLUID |  | STATE | RANGE (kg/cm2g) |  |  | PRESSURE (kg/cm2g) |  |  | TEMP. (deg C) |  |  | ACCESSORIES |  |  |  | REMARK
[30 empty] | DESIGN PRESS. |  | DESIGN TEMP. |  | REMARK
[16 empty] | MAX | NOR |  | MAX | NOR | [6 empty] | TYPE | SPEC. | RATING | INT. | EXT. | HOT | COLD
1 | PG-1211 | HP TO EA-211 |  |  |  |  | D-1-IPA-2001 |  |  | SH |  | S | 0 ~ 50 |  |  |  | 28 |  |  | 231 |  | S |  | LF | SF |  | 1/2" NPT | BA59C | NPT | 46 | 0 | 400 | 0 |  | O
2 | PG-1801 | HP FROM BL |  |  |  |  | D-1-IPA-8011 |  |  | HS |  | S | 0 ~ 70 |  |  |  | 40 |  |  | 370 |  | S |  | LF | SF |  | 1/2" NPT | BA59C | NPT | 45 | F.V | 425 | 0 |  | O
3 | PG-1805 | MSS TO BG-803 |  |  |  |  | D-1-IPA-8011 |  |  | MSS |  | S | 0 ~ 30 |  |  |  | 14.6 |  |  | 226 |  | S |  | LF |  |  | 1" 150# RF | BB57C | NPT | 18 | 0 | 350 | 0 |  | O
4 | PG-1806 | MSS TO BG-803 |  |  |  |  | D-1-IPA-8011 |  |  | MSS |  | S | 0 ~ 30 |  |  |  | 14.6 |  |  | 226 |  | S |  | LF |  |  | 1/2" NPT | BB57C | NPT | 18 | 0 | 350 | 0 |  | O
5 | PG-1811 | HP TO BG-801 |  |  |  |  | D-1-IPA-8011 |  |  | HS |  | S | 0 ~ 70 |  |  |  | 40 |  |  | 370 |  | S |  | LF | SF |  | 1/2" NPT | BA59C | NPT | 45 | F.V | 425 | 0 |  | O
6 | PG-1812 | HP TO BG-802 |  |  |  |  | D-1-IPA-8011 |  |  | HS |  | S | 0 ~ 70 |  |  |  | 40 |  |  | 370 |  | S |  | LF | SF |  | 1/2" NPT | BA59C | NPT | 45 | F.V | 425 | 0 |  | O

## Sheet2
[12 empty] | PRESSURE (kg/cm2g) |  | TEMP. (deg C) |  | ACCESSORIES |  |  | REMARK
| NO | TAG NO | INST TYPE | SERVICE | P&ID NO | FLUID | STATE | RANGE (kg/cm2g) | [14 empty] | DESIGN PRESS. |  | DESIGN TEMP.
[12 empty] | MAX | NOR | MAX | NOR |  |  |  |  | TYPE | SPEC. | RATING | INT. | EXT. | HOT | COLD | 1.5 | 1.3 | NOR/RANGE
[27 empty] | NOR x 1.5 | ra x 1.3
| 1 | PG-1134 | DIAPHRAGM 150# | FA-134 EQUIP. | D-1-IPA-0004-2 | PRG | L | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.5 |  | 25 |  |  | LF |  | 1" 150# RF | FF55C | 150# RF | 8 | 0 | 60 | 0 | 0.75 | 1.3 | 0.75
| 2 | PG-1135 | DIAPHRAGM 150# | FA-134 EQUIP. | D-1-IPA-0004-2 | PRG | L | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.5 |  | 25 |  |  | LF |  | 1" 150# RF | FF55C | 150# RF | 8 | 0 | 60 | 0 | 0.75 | 1.3 | 0.75
| 3 | PG-1136 | DIAPHRAGM 150# | GA-134 DISCH. | D-1-IPA-0004-2 | PDI | L | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2.47 |  | 25 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10.7 | 0 | 60 | 0 | 3.705 | 6.5 | 0.494
| 4 | PG-1137 | DIAPHRAGM 150# | FA-134 EQUIP. | D-1-IPA-0004-2 | PDI | L | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.5 |  | 25 |  | D | LF |  | 1" 150# RF | - | 150# RF | 8 | 0 | 60 | 0 | 0.75 | 1.3 | 0.75
| 5 | PG-1138 | DIAPHRAGM 150# | GA-112A/B SUCTION | D-1-IPA-1001 | PCS | L | 0 ~ 40 | -1 | 1 | -1 ~ 1 | 1.58 | ATM |  | 20 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 0.5 | 0 | 120 | 0 | #VALUE! | 1.3 | #VALUE!
| 6 | PG-1155A | DIAPHRAGM 150# | GA-155A DISCH. | D-1-IPA-1003 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.39 |  | 82 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.9 | 0 | 120 | 0 | 8.085 | 13 | 0.539
| 7 | PG-1155B | DIAPHRAGM 150# | GA-155B DISCH. | D-1-IPA-1003 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.39 |  | 82 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.9 | 0 | 120 | 0 | 8.085 | 13 | 0.539
| 8 | PG-1163A | DIAPHRAGM 150# | GA-163A DISCH. | D-1-IPA-1003 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.71 |  | 78 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.6 | 0 | 120 | 0 | 8.565 | 13 | 0.571
| 9 | PG-1163B | DIAPHRAGM 150# | GA-163B DISCH. | D-1-IPA-1003 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.71 |  | 78 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.6 | 0 | 120 | 0 | 8.565 | 13 | 0.571
| 10 | PG-1170-1 | DIAPHRAGM 150# | FC-170 EQUIP. | D-1-IPA-1003 | PRG | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.1 |  | 38 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 2 | 0 | 120 | 0 | 0.15 | 1.3 | 0.55
| 11 | PG-1221 | DIAPHRAGM 150# | LP TO EA-221 | D-1-IPA-2002 | SL | S | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2.5 |  | 138.2 |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 6 | 0 | 200 | 0 | 3.75 | 6.5 | 0.5
| 12 | PG-1313A | DIAPHRAGM 150# | GA-313A DISCH. | D-1-IPA-3001 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.73 |  | 97 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.8 | 0 | 150 | 0 | 8.595 | 13 | 0.573
| 13 | PG-1313B | DIAPHRAGM 150# | GA-313B DISCH. | D-1-IPA-3001 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.73 |  | 97 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.8 | 0 | 150 | 0 | 8.595 | 13 | 0.573
| 14 | PG-1315A | DIAPHRAGM 150# | GA-315A DISCH. | D-1-IPA-3001 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.25 |  | 45 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.4 | 0 | 120 | 0 | 7.875 | 13 | 0.525
| 15 | PG-1315B | DIAPHRAGM 150# | GA-315B DISCH. | D-1-IPA-3001 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.25 |  | 45 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.4 | 0 | 120 | 0 | 7.875 | 13 | 0.525
| 16 | PG-1317A | DIAPHRAGM 150# | GA-317A DISCH. | D-1-IPA-3001 | PE | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 3.57 |  | 45 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.7 | 0 | 120 | 0 | 5.355 | 13 | 0.357
| 17 | PG-1317B | DIAPHRAGM 150# | GA-317B DISCH. | D-1-IPA-3001 | PE | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 3.57 |  | 45 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.7 | 0 | 120 | 0 | 5.355 | 13 | 0.357
| 18 | PG-1318 | DIAPHRAGM 150# | N2 TO FLARE | D-1-IPA-3001 | NL | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.01 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 0.015 | 1.3 | 0.505 | P&ID 상 위치 PRV(SET:0.01kg/cm2.g) 후단
| 19 | PG-1319 | DIAPHRAGM 150# | DA-310 EQUIP. | D-1-IPA-3001 | PDI | L | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.25 |  | 97 |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 150 | 0 | 0.375 | 1.3 | 0.625
| 20 | PG-1323A | DIAPHRAGM 150# | GA-323A DISCH. | D-1-IPA-3002 | PWW | L | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 8.23 |  | 133 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10.6 | 0 | 180 | 0 | 12.345 | 19.5 | 0.54867
| 21 | PG-1323B | DIAPHRAGM 150# | GA-323B DISCH. | D-1-IPA-3002 | PWW | L | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 8.23 |  | 133 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10.6 | 0 | 180 | 0 | 12.345 | 19.5 | 0.54867
| 22 | PG-1325A | DIAPHRAGM 150# | GA-325A DISCH. | D-1-IPA-3003 | PDI | L | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7.97 |  | 50 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10.2 | 0 | 140 | 0 | 11.955 | 19.5 | 0.53133
| 23 | PG-1325B | DIAPHRAGM 150# | GA-325B DISCH. | D-1-IPA-3003 | PDI | L | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7.97 |  | 50 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10.2 | 0 | 140 | 0 | 11.955 | 19.5 | 0.53133
공정질의 | 24 | PG-1327 | DIAPHRAGM 150# | N2 TO FLARE | D-1-IPA-3003 | NL | V | 0 ~ 5 | -1 | 1 | -1 ~ 1 |  | 0.01 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 0.015 | 1.3 | 0.505 | P&ID 상 위치 PRV(SET:0.01kg/cm2.g) 후단
| 25 | PG-1328 | DIAPHRAGM 150# | FA-325 EQUIP. | D-1-IPA-3003 | PDI | L | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.02 |  | 50 |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 140 | 0 | 0.03 | 1.3 | 0.51
| 26 | PG-1330A | DIAPHRAGM 150# | GA-330A DISCH. | D-1-IPA-3002 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.37 |  | 111 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.5 | 0 | 180 | 0 | 9.555 | 13 | 0.637
| 27 | PG-1330B | DIAPHRAGM 150# | GA-330B DISCH. | D-1-IPA-3002 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.37 |  | 111 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.5 | 0 | 180 | 0 | 9.555 | 13 | 0.637
| 28 | PG-1331 | DIAPHRAGM 150# | LP TO EA-331 | D-1-IPA-3002 | SL | S | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2.5 |  | 138.2 |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 6 | 0 | 200 | 0 | 3.75 | 6.5 | 0.5
| 29 | PG-1331A | DIAPHRAGM 150# | GA-333A DISCH. | D-1-IPA-3002 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.8 |  | 117 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.9 | 0 | 160 | 0 | 10.2 | 13 | 0.68
| 30 | PG-1331B | DIAPHRAGM 150# | GA-333B DISCH. | D-1-IPA-3002 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.8 |  | 117 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.9 | 0 | 160 | 0 | 10.2 | 13 | 0.68
| 31 | PG-1345 | DIAPHRAGM 150# | N2 TO FLARE | D-1-IPA-3003 | NL | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.005 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 0.0075 | 1.3 | 0.5025 | P&ID 상 위치 PRV(SET:0.005kg/cm2.g) 후단
| 32 | PG-1351 | DIAPHRAGM 150# | DA-350 EQUIP. | D-1-IPA-3004 | PIP | L | -1 ~ 1 | -1 | 2 | -1 ~ 2 |  | 0.4 |  | 90 |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 120 | 0 | 0.6 | 2.6 | 0.46667
| 33 | PG-1351A | DIAPHRAGM 150# | EA-351A EQUIP. | D-1-IPA-3004 | PDI | V | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 1.6 |  | 105.6 |  | D | LF |  | 1" 150# RF | - | 150# RF | 4 | 0 | 150 | 0 | 2.4 | 6.5 | 0.32
| 34 | PG-1351B | DIAPHRAGM 150# | EA-351B EQUIP. | D-1-IPA-3004 | PDI | V | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 1.6 |  | 105.6 |  | D | LF |  | 1" 150# RF | - | 150# RF | 4 | 0 | 150 | 0 | 2.4 | 6.5 | 0.32
| 35 | PG-1353A | DIAPHRAGM 150# | GA-353A DISCH. | D-1-IPA-3004 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.37 |  | 90 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6 | 0 | 120 | 0 | 6.555 | 13 | 0.437
| 36 | PG-1353B | DIAPHRAGM 150# | GA-353B DISCH. | D-1-IPA-3004 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.37 |  | 90 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6 | 0 | 120 | 0 | 6.555 | 13 | 0.437
| 37 | PG-1355A | DIAPHRAGM 150# | GA-355A DISCH. | D-1-IPA-3004 | PS | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.73 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.1 | 0 | 120 | 0 | 7.095 | 13 | 0.473
| 38 | PG-1355B | DIAPHRAGM 150# | GA-355B DISCH. | D-1-IPA-3004 | PS | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.73 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7.1 | 0 | 120 | 0 | 7.095 | 13 | 0.473
| 39 | PG-1356A | DIAPHRAGM 150# | GA-356A DISCH. | D-1-IPA-3004 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 3.44 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.7 | 0 | 120 | 0 | 5.16 | 13 | 0.344
| 40 | PG-1356B | DIAPHRAGM 150# | GA-356B DISCH. | D-1-IPA-3004 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 3.44 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.7 | 0 | 120 | 0 | 5.16 | 13 | 0.344
| 41 | PG-1357 | DIAPHRAGM 150# | DA-350 EQUIP. | D-1-IPA-3004 | PDI / PIP | M | -1 ~ 1 | -1 | 2 | -1 ~ 2 |  | 0.4 |  | 90 |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 120 | 0 | 0.6 | 2.6 | 0.46667
| 42 | PG-1358 | DIAPHRAGM 150# | N2 TO FLARE | D-1-IPA-3005 | NL | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.01 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 0.015 | 1.3 | 0.505 | P&ID 상 위치 PRV(SET:0.01kg/cm2.g) 후단
| 43 | PG-1361 | DIAPHRAGM 150# | LP TO EA-361 | D-1-IPA-3006 | SL | S | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2.5 |  | 138.2 |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 6 | 0 | 200 | 0 | 3.75 | 6.5 | 0.5
| 44 | PG-1362 | DIAPHRAGM 150# | EA-361 EQUIP. | D-1-IPA-3006 | SL | S | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2.5 |  | 138.2 |  | D | LF |  | 1" 150# RF | - | 150# RF | 6 | 0 | 200 | 0 | 3.75 | 6.5 | 0.5
| 45 | PG-1363A | DIAPHRAGM 150# | GA-363A DISCH. | D-1-IPA-3006 | PPW | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.49 |  | 108 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.4 | 0 | 140 | 0 | 9.735 | 13 | 0.649
| 46 | PG-1363B | DIAPHRAGM 150# | GA-363B DISCH. | D-1-IPA-3006 | PPW | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.49 |  | 108 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.4 | 0 | 140 | 0 | 9.735 | 13 | 0.649
| 47 | PG-1364 | DIAPHRAGM 150# | DA-360 EQUIP. | D-1-IPA-3006 | PPW | L | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.3 |  | 108 |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 140 | 0 | 0.45 | 1.3 | 0.65
| 48 | PG-1365A | DIAPHRAGM 150# | GA-365A DISCH. | D-1-IPA-3006 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.76 |  | 74 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.9 | 0 | 120 | 0 | 7.14 | 13 | 0.476
| 49 | PG-1365B | DIAPHRAGM 150# | GA-365B DISCH. | D-1-IPA-3006 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.76 |  | 74 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.9 | 0 | 120 | 0 | 7.14 | 13 | 0.476
| 50 | PG-1367 | DIAPHRAGM 150# | EA-365 EQUIP. | D-1-IPA-3006 | PDI | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.05 |  | 80.3 |  | D | LF |  | 1" 150# RF | - | 150# RF | 6 | 0 | 120 | 0 | 0.075 | 1.3 | 0.525
| 51 | PG-1371 | DIAPHRAGM 150# | FA-371 EQUIP. | D-1-IPA-3007 | VA | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.03 / -0.02 |  | 50 |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 | #VALUE! | 1.3 | #VALUE!
| 52 | PG-1371A | DIAPHRAGM 150# | GA-371A DISCH. | D-1-IPA-3007 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.19 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.1 | 0 | 120 | 0 | 7.785 | 13 | 0.519
| 53 | PG-1371B | DIAPHRAGM 150# | GA-371B DISCH. | D-1-IPA-3007 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.19 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.1 | 0 | 120 | 0 | 7.785 | 13 | 0.519
| 54 | PG-1372A | DIAPHRAGM 150# | FA-373A EQUIP. | D-1-IPA-3007 | GFL | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.03 / -0.02 |  | 40 |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 | #VALUE! | 1.3 | #VALUE!
| 55 | PG-1372B | DIAPHRAGM 150# | FA-373B EQUIP. | D-1-IPA-3007 | GFL | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.03 / -0.02 |  | 40 |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 | #VALUE! | 1.3 | #VALUE!
| 56 | PG-1372C | DIAPHRAGM 150# | FA-373C EQUIP. | D-1-IPA-3007 | GFL | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.03 / -0.02 |  | 40 |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 | #VALUE! | 1.3 | #VALUE!
| 57 | PG-1373A | DIAPHRAGM 150# | GA-373A DISCH. | D-1-IPA-3007 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.66 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.6 | 0 | 120 | 0 | 6.99 | 13 | 0.466
| 58 | PG-1373B | DIAPHRAGM 150# | GA-373B DISCH. | D-1-IPA-3007 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.66 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.6 | 0 | 120 | 0 | 6.99 | 13 | 0.466
| 59 | PG-1374 | DIAPHRAGM 150# | N2 TO FLARE | D-1-IPA-3007 | NL | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.01 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 0.015 | 1.3 | 0.505 | P&ID 상 위치 PRV(SET:0.01kg/cm2.g) 후단
| 60 | PG-1374A | DIAPHRAGM 150# | GA-374A DISCH. | D-1-IPA-3007 | PDI | L | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2.1 |  | 70 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 3.5 | 0 | 150 | 0 | 3.15 | 6.5 | 0.42
| 61 | PG-1374B | DIAPHRAGM 150# | GA-374B DISCH. | D-1-IPA-3007 | PDI | L | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2.1 |  | 70 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 3.5 | 0 | 150 | 0 | 3.15 | 6.5 | 0.42
| 62 | PG-1375 | DIAPHRAGM 150# | FD-373A/B DISCH. | D-1-IPA-3007 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.66 |  | 40 |  |  | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.6 | 0 | 100 | 0 | 6.99 | 13 | 0.466
| 63 | PG-1376 | DIAPHRAGM 150# | N2 TO FLARE | D-1-IPA-3007 | NL | V | 0 ~ 15 | 0 | 5 | 0 ~ 5 |  | 2 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 3 | 6.5 | 0.4 | P&ID 상 위치 PRV(SET:2kg/cm2.g) 후단
| 64 | PG-1381 | DIAPHRAGM 150# | N2 TO FA-371 | D-1-IPA-3007 | NL | V | 0 ~ 5 | 0 | 10 | 0 ~ 10 |  | 2 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 3 | 13 | 0.2
| 65 | PG-1401 | DIAPHRAGM 150# | IPA TO LOADING SYSTEM | D-1-IPA-4004 | PIP | L | 0 ~ 5 | 0 | 3 | 0 ~ 3 |  | 1.54 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 3.41 | 0 | 120 | 0 | 2.31 | 3.9 | 0.51333
| 66 | PG-1402 | DIAPHRAGM 150# | N2 TO LOADING SYSTEM | D-1-IPA-4004 | NL | V | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7 |  | AMB |  | D | LF | CONN. : 1/2" 150# RF | 1" 150# RF | BB55C1 | 150# RF | 10 | 0 | 60 | 0 | 10.5 | 19.5 | 0.46667
| 67 | PG-1403 | DIAPHRAGM 150# | IPA TO LOADING SYSTEM | D-1-IPA-4004 | PIP | L | 0 ~ 5 | 0 | 3 | 0 ~ 3 |  | 1.54 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 3.41 | 0 | 120 | 0 | 2.31 | 3.9 | 0.51333
| 68 | PG-1404 | DIAPHRAGM 150# | N2 TO LOADING SYSTEM | D-1-IPA-4004 | NL | V | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7 |  | AMB |  | D | LF | CONN. : 1/2" 150# RF | 1" 150# RF | BB55C1 | 150# RF | 10 | 0 | 60 | 0 | 10.5 | 19.5 | 0.46667
| 69 | PG-1405 | DIAPHRAGM 150# | IPE TO LOADING SYSTEM | D-1-IPA-4004 | PDI | L | 0 ~5 | 0 | 5 | 0 ~ 5 |  | 2.83 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.24 | 0 | 120 | 0 | 4.245 | 6.5 | 0.566
| 70 | PG-1406 | DIAPHRAGM 150# | N2 TO LOADING SYSTEM | D-1-IPA-4004 | NL | V | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7 |  | AMB |  | D | LF | CONN. : 1/2" 150# RF | 1" 150# RF | BB55C1 | 150# RF | 10 | 0 | 60 | 0 | 10.5 | 19.5 | 0.46667
| 71 | PG-1420A | DIAPHRAGM 150# | GA-420A DISCH. | D-1-IPA-4002 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.25 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.52 | 0 | 120 | 0 | 6.375 | 13 | 0.425
| 72 | PG-1420B | DIAPHRAGM 150# | GA-420B DISCH. | D-1-IPA-4002 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.25 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.52 | 0 | 120 | 0 | 6.375 | 13 | 0.425
| 73 | PG-1422 | DIAPHRAGM 150# | N2 TO FB-420A/B | D-1-IPA-4002 | NL | V | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 3 | 6.5 | 0.4
| 74 | PG-1424 | DIAPHRAGM 150# | FD-420A INLET | D-1-IPA-4002 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.25 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.52 | 0 | 120 | 0 | 6.375 | 13 | 0.425
| 75 | PG-1425 | DIAPHRAGM 150# | FD-420A OUTLET | D-1-IPA-4002 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.1 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.5 | 0 | 120 | 0 | 6.15 | 13 | 0.41
| 76 | PG-1426 | DIAPHRAGM 150# | FD-420B INLET | D-1-IPA-4002 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.25 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.52 | 0 | 120 | 0 | 6.375 | 13 | 0.425
| 77 | PG-1427 | DIAPHRAGM 150# | FD-420B OUTLET | D-1-IPA-4002 | PIP | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.1 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 6.5 | 0 | 120 | 0 | 6.15 | 13 | 0.41
| 78 | PG-1430 | DIAPHRAGM 150# | GA-430 DISCH. | D-1-IPA-4003 | PPA | L | 0 ~ 5 | 0 | 3 | 0 ~ 3 |  | 1.2 |  | AMB |  | D | LF |  | 1" 150# RF | FF55C | 150# RF | 2.1 | 0 | 120 | 0 | 1.8 | 3.9 | 0.4
| 79 | PG-1431 | DIAPHRAGM 150# | GA-431 DISCH. | D-1-IPA-4003 | PS | L | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2.19 |  | 20 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.9 | 0 | 120 | 0 | 3.285 | 6.5 | 0.438
| 80 | PG-1435 | DIAPHRAGM 150# | NS FROM GA-982XC | D-1-IPA-4003 | NS | L | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7.7 |  | 40 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 10 | 0 | 120 | 0 | 11.55 | 19.5 | 0.51333
| 81 | PG-1781 | DIAPHRAGM 150# | FA-781A EQUIP. | D-1-IPA-7001 | PDI | V | 0 ~ 40 | -1 | 1 | -1 ~ 1 |  | ATM |  | AMB |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 | #VALUE! | 1.3 | #VALUE!
| 82 | PG-1782 | DIAPHRAGM 150# | FA-781B EQUIP. | D-1-IPA-7001 | PDI | V | 0 ~ 40 | -1 | 1 | -1 ~ 1 |  | ATM |  | AMB |  | D | LF |  | 1" 150# RF | - | 150# RF | 0.5 | -0.1 | 120 | 0 | #VALUE! | 1.3 | #VALUE!
| 83 | PG-1783 | DIAPHRAGM 150# | GA-781A DISCH. | D-1-IPA-7001 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.57 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.58 | 0 | 120 | 0 | 6.855 | 13 | 0.457
| 84 | PG-1784 | DIAPHRAGM 150# | GA-781B DISCH. | D-1-IPA-7001 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4.57 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 5.58 | 0 | 120 | 0 | 6.855 | 13 | 0.457
| 85 | PG-1808 | DIAPHRAGM 150# | SL TO CONSUMERS | D-1-IPA-8011 | SL | S | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 3.5 |  | 148 |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 6 | F.V | 200 | 0 | 5.25 | 13 | 0.35
| 86 | PG-1815A | DIAPHRAGM 150# | GA-815A DISCH. | D-1-IPA-8011 | CL | L | 0 ~ 20 | 0 | 20 | 0 ~ 20 |  | 11.3 |  | 100 |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 14.3 | 0 | 150 | 0 | 16.95 | 26 | 0.565
| 87 | PG-1815B | DIAPHRAGM 150# | GA-815B DISCH. | D-1-IPA-8011 | CL | L | 0 ~ 20 | 0 | 20 | 0 ~ 20 |  | 11.3 |  | 100 |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 14.3 | 0 | 150 | 0 | 16.95 | 26 | 0.565
| 88 | PG-1820A | DIAPHRAGM 150# | GA-820A DISCH. | D-1-IPA-8012 | WCS | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.5 |  | 32 |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 8.5 | 0 | AMB | 0 | 9.75 | 13 | 0.65
| 89 | PG-1820B | DIAPHRAGM 150# | GA-820B DISCH. | D-1-IPA-8012 | WCS | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.5 |  | 32 |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 8.5 | 0 | AMB | 0 | 9.75 | 13 | 0.65
| 90 | PG-1821 | DIAPHRAGM 150# | WCR TO EF-820 | D-1-IPA-8012 | WCR | L | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2 |  | 42 |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 7 | 0 | 50 | 0 | 3 | 6.5 | 0.4
| 91 | PG-1830A | DIAPHRAGM 150# | GA-830A DISCH. | D-1-IPA-8013 | WT | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.4 |  | 120 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 9.5 | 0 | 180 | 0 | 9.6 | 13 | 0.64
| 92 | PG-1830B | DIAPHRAGM 150# | GA-830B DISCH. | D-1-IPA-8013 | WT | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.4 |  | 120 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 9.5 | 0 | 180 | 0 | 9.6 | 13 | 0.64
| 93 | PG-1831 | DIAPHRAGM 150# | FA-830X EQUIP. | D-1-IPA-8013 | WT | L | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.25 | 120 | 105 |  | D | LF |  | 1" 150# RF | - | 150# RF | 4 | F.V | 180 | 0 | 0.375 | 1.3 | 0.625
| 94 | PG-1835A | DIAPHRAGM 150# | GA-835A DISCH. | D-1-IPA-8013 | WT | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.7 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.7 | 0 | 150 | 0 | 10.05 | 13 | 0.67
| 95 | PG-1835B | DIAPHRAGM 150# | GA-835B DISCH. | D-1-IPA-8013 | WT | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 6.7 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 8.7 | 0 | 150 | 0 | 10.05 | 13 | 0.67
| 96 | PG-1838 | DIAPHRAGM 150# | DA-830 EQUIP. | D-1-IPA-8013 | WT | L | -1 ~ 1 | -1 | 1 | -1 ~ 1 | 1 | 0.25 |  | 105 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 4 | F.V | 180 | 0 | 0.375 | 1.3 | 0.625
| 97 | PG-1839 | DIAPHRAGM 150# | WT FROM BL | D-1-IPA-8013 | WT | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 7 | 0 | AMB | 0 | 7.5 | 13 | 0.5
| 98 | PG-1842 | DIAPHRAGM 150# | AI FROM BL | D-1-IPA-8014 | AI | V | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7 |  | AMB |  | D | LF |  | 1" 150# RF | BC55C | 150# RF | 10 | 0 | 60 | 0 | 10.5 | 19.5 | 0.46667
| 99 | PG-1843 | DIAPHRAGM 150# | FA-841 TO AI HEADER | D-1-IPA-8014 | AI | V | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7 |  | AMB |  | D | LF |  | 1" 150# RF | BC55C | 150# RF | 10 | 0 | 60 | 0 | 10.5 | 19.5 | 0.46667
| 100 | PG-1844 | DIAPHRAGM 150# | FA-841 EQUIP. | D-1-IPA-8014 | AI | V | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7 |  | AMB |  | D | LF |  | 1" 150# RF | - | 150# RF | 10 | 0 | 60 | 0 | 10.5 | 19.5 | 0.46667
| 101 | PG-1847 | DIAPHRAGM 150# | NL TO NL HEADER | D-1-IPA-8014 | NL | V | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 10.5 | 19.5 | 0.46667
| 102 | PG-1849 | DIAPHRAGM 150# | NL FROM BL | D-1-IPA-8014 | NL | V | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 10.5 | 19.5 | 0.46667
| 103 | PG-1850A | DIAPHRAGM 150# | GA-840A DISCH. | D-1-IPA-8024 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.5 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 7.9 | 0 | 70 | 0 | 8.25 | 13 | 0.55
| 104 | PG-1850B | DIAPHRAGM 150# | GA-840B DISCH. | D-1-IPA-8024 | PDI | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 5.5 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 7.9 | 0 | 70 | 0 | 8.25 | 13 | 0.55
| 105 | PG-1851A | DIAPHRAGM 150# | GA-850A DISCH. | D-1-IPA-8024 | WW | L | 0 ~ 5 | 0 | 3 | 0 ~ 3 |  | 1.55 |  | 30 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 2.7 | 0 | 70 | 0 | 2.325 | 3.9 | 0.51667
| 106 | PG-1851B | DIAPHRAGM 150# | GA-850B DISCH. | D-1-IPA-8024 | WW | L | 0 ~ 5 | 0 | 3 | 0 ~ 3 |  | 1.55 |  | 30 |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 2.7 | 0 | 70 | 0 | 2.325 | 3.9 | 0.51667
| 107 | PG-1860 | DIAPHRAGM 150# | FA-860 EQUIP. | D-1-IPA-8023-1 | VFL | V | -1 ~ 1 | -1 | 1 | -1 ~ 1 |  | 0.3 |  | AMB |  | D | LF |  | 1" 150# RF | - | 150# RF | 2 | 0 | 300 | 0 | 0.45 | 1.3 | 0.65
| 108 | PG-1860A | DIAPHRAGM 150# | GA-860A DISCH. | D-1-IPA-8023-1 | PDI | L | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2.18 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 4 | 0 | 300 | 0 | 3.27 | 6.5 | 0.436
| 109 | PG-1860B | DIAPHRAGM 150# | GA-860B DISCH. | D-1-IPA-8023-1 | PDI | L | 0 ~ 5 | 0 | 5 | 0 ~ 5 |  | 2.18 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C1 | 150# RF | 4 | 0 | 300 | 0 | 3.27 | 6.5 | 0.436
| 110 | PG-1870 | DIAPHRAGM 150# | AP FROM BL | D-1-IPA-8015 | AP | V | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 7 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 10 | 0 | 60 | 0 | 10.5 | 19.5 | 0.46667
| 111 | PG-1880 | DIAPHRAGM 150# | WF FROM BL | D-1-IPA-8022 | WR | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 |  | 4 |  | AMB |  | D | LF |  | 1" 150# RF | BB55C | 150# RF | 7 | 0 | 60 | 0 | 6 | 13 | 0.4
| 112 | PG-1900A | DIAPHRAGM 150# | GA-900A DISCH. | D-1-IPA-8024 | WW | L | HOLD | HOLD |  | HOLD ~ |  | HOLD |  | HOLD |  | D | LF | HOLD | 1" 150# RF | -- | 150# RF | -- | -- | -- | -- | #VALUE! | 0 | #VALUE!
| 113 | PG-1900B | DIAPHRAGM 150# | GA-900B DISCH. | D-1-IPA-8024 | WW | L | HOLD | HOLD |  | HOLD ~ |  | HOLD |  | HOLD |  | D | LF | HOLD | 1" 150# RF | -- | 150# RF | -- | -- | -- | -- | #VALUE! | 0 | #VALUE!
| 1 | PG-1215A | DIAPHRAGM 300# | GA-215A DISCH. | D-1-IPA-2001 | PPR | L | 0 ~ 40 | 0 | 40 | 0 ~ 40 |  | 22.9 |  | 40 |  | D | LF |  | 1" 300# RF | BB57C1 | 300# RF | 31.7 | 0 | 120 | 0 | 34.35 | 52 | 0.5725
| 2 | PG-1215B | DIAPHRAGM 300# | GA-215B DISCH. | D-1-IPA-2001 | PPR | L | 0 ~ 40 | 0 | 40 | 0 ~ 40 |  | 22.9 |  | 40 |  | D | LF |  | 1" 300# RF | BB57C1 | 300# RF | 31.7 | 0 | 120 | 0 | 34.35 | 52 | 0.5725
| 3 | PG-1220 | DIAPHRAGM 300# | DA-220 EQUIP. | D-1-IPA-2002 | PHE / PPR | M | 0 ~ 30 | 0 | 30 | 0 ~ 30 |  | 19.6 |  | 52 |  |  | LF |  | 1" 300# RF | - | 300# RF | 25 | 0 | 120 | 0 | 29.4 | 39 | 0.65333
| 4 | PG-1227 | DIAPHRAGM 300# | GA-226A/B SUCTION | D-1-IPA-2002 | PPR | L | 0 ~ 30 | 0 | 30 | 0 ~ 30 |  | 19.32 |  | 47.2 |  |  | LF |  | 1" 300# RF | BB57C1 | 300# RF | 25 | 0 | 120 | -46 | 28.98 | 39 | 0.644
| 5 | PG-1311 | DIAPHRAGM 300# | MP STEAM TO EA-311 | D-1-IPA-3001 | SM | S | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 8 |  | 174.5 |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | 0 | 250 | 0 | 12 | 19.5 | 0.53333
| 6 | PG-1321 | DIAPHRAGM 300# | MP STEAM TO EA-321 | D-1-IPA-3002 | SM | S | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 8 |  | 174.5 |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | 0 | 250 | 0 | 12 | 19.5 | 0.53333
| 7 | PG-1352 | DIAPHRAGM 300# | EA-352 EQUIP. | D-1-IPA-3004 | SM | S | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 8 |  | 174.5 |  |  | LF |  | 1" 300# RF | - | 300# RF | 12 | 0 | 250 | 0 | 12 | 19.5 | 0.53333
| 8 | PG-1354 | DIAPHRAGM 300# | EA-352 EQUIP. | D-1-IPA-3004 | SM | S | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 8 |  | 174.5 |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | 0 | 250 | 0 | 12 | 19.5 | 0.53333
| 9 | PG-1401A | DIAPHRAGM 300# | GA-410A DISCH. | B-1-9112 | PPR | L | 0 ~ 30 | 0 | 30 | 0 ~ 30 |  | 18 | 25 | 35 |  | D | LF |  | 1" 300# RF | BA8A | 300# RF | 25 | 0 | 65 | -15 | 27 | 39 | 0.6
| 10 | PG-1401B | DIAPHRAGM 300# | GA-410B DISCH. | B-1-9112 | PPR | L | 0 ~ 30 | 0 | 30 | 0 ~ 30 |  | 18 | 25 | 35 |  | D | LF |  | 1" 300# RF | BA8A | 300# RF | 25 | 0 | 65 | -15 | 27 | 39 | 0.6
| 11 | PG-1412 | DIAPHRAGM 300# | PPR FROM GA-410A/B | D-1-IPA-4001 | PPR | L | 0 ~ 30 | 0 | 30 | 0 ~ 30 |  | 18.01 | 35 | 25 |  |  | LF |  | 1" 300# RF | BA8A | 300# RF | 25 | 0 | 65 | -15 | 27.015 | 39 | 0.60033
| 12 | PG-1415 | DIAPHRAGM 300# | FA-415 EQUIP. | D-1-IPA-4001 | VFL | V | 0 ~ 30 | 0 | 30 | 0 ~ 30 |  | 15.2 |  | AMB |  |  | LF |  | 1" 300# RF | - | 300# RF | 25 | 0 | 65 | -15 | 22.8 | 39 | 0.50667
| 13 | PG-1418A | DIAPHRAGM 300# | GA-415A SUCTION | D-1-IPA-4001 | PPF | L | 0 ~ 30 | 0 | 30 | 0 ~ 30 |  | 15.4 | 35 | 25 |  |  | LF |  | 1" 300# RF | BB57C1 | 300# RF | 25 | 0 | 65 | -15 | 23.1 | 39 | 0.51333
| 14 | PG-1418B | DIAPHRAGM 300# | GA-415B SUCTION | D-1-IPA-4001 | PPF | L | 0 ~ 30 | 0 | 30 | 0 ~ 30 |  | 15.4 | 35 | 25 |  |  | LF |  | 1" 300# RF | BB57C1 | 300# RF | 25 | 0 | 65 | -15 | 23.1 | 39 | 0.51333
| 15 | PG-1807 | DIAPHRAGM 300# | SM TO CONSUMERS | D-1-IPA-8011 | SM | S | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 9.8 |  | 198 |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | F.V | 250 | 0 | 14.7 | 19.5 | 0.65333
| 16 | PG-1833 | DIAPHRAGM 300# | DA-830 TO EA-830X INLET | D-1-IPA-8013 | CM | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 | 8 | 2.5 | 193.5 | 143.5 |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | F.V | 250 | 0 | 3.75 | 13 | 0.25
| 17 | PG-1834 | DIAPHRAGM 300# | EA-830X OUTLET TO CM | D-1-IPA-8013 | CM | L | 0 ~ 10 | 0 | 10 | 0 ~ 10 | 8 | 2.5 | 193.5 | 143.5 |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 12 | F.V | 250 | 0 | 3.75 | 13 | 0.25
| 18 | PG-1841 | DIAPHRAGM 300# | FA-845A/B TO FA-841 | D-1-IPA-8014 | AI | V | 0 ~ 15 | 0 | 15 | 0 ~ 15 |  | 10 |  | AMB |  |  | LF |  | 1" 300# RF | BB57C | 300# RF | 13 | 0 | 65 | 0 | 15 | 19.5 | 0.66667
| 1 | PG-1110A | DIAPHRAGM 600# | GA-110A TO DC-110 | D-1-IPA-1001 | PPA | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 50.34 |  | 20 |  | D | LF |  | 1" 600# RF | FF59C | 600# RF | 60 | 0 | 120 | 0 | 75.51 | 104 | 0.62925
| 2 | PG-1110B | DIAPHRAGM 600# | GA-110A TO DC-110 | D-1-IPA-1001 | PPA | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 50.34 |  | 20 |  | D | LF |  | 1" 600# RF | FF59C | 600# RF | 60 | 0 | 120 | 0 | 75.51 | 104 | 0.62925
| 3 | PG-1112A | DIAPHRAGM 600# | GA-112A DISCH. | D-1-IPA-1001 | PCS | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 51.2 |  | 20 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 60 | 0 | 120 | 0 | 76.8 | 104 | 0.64
| 4 | PG-1112B | DIAPHRAGM 600# | GA-112B DISCH. | D-1-IPA-1001 | PCS | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 51.2 |  | 20 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 60 | 0 | 120 | 0 | 76.8 | 104 | 0.64
| 5 | PG-1113 | DIAPHRAGM 600# | GA-113 DISCH. | D-1-IPA-1001 | PCS | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 51.3 |  | 20 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 60 | 0 | 120 | 0 | 76.95 | 104 | 0.64125
| 6 | PG-1115A | DIAPHRAGM 600# | GA-115A DISCH. | D-1-IPA-1001 | PPW | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 49.8 |  | 103 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 61.1 | 0 | 160 | 0 | 74.7 | 104 | 0.6225
| 7 | PG-1115B | DIAPHRAGM 600# | GA-115B DISCH. | D-1-IPA-1001 | PPW | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 49.8 |  | 103 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 61.1 | 0 | 160 | 0 | 74.7 | 104 | 0.6225
| 8 | PG-1140 | DIAPHRAGM 600# | GB-120 DISCH. | D-1-IPA-1002 | PRG | V | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 47 |  | 125 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 55 | 0 | 150 | 0 | 70.5 | 104 | 0.5875
| 9 | PG-1141 | DIAPHRAGM 600# | FA-121 EQUIP. | D-1-IPA-1002 | PRG | V | 0 ~ 70 | 0 | 70 | 0 ~ 70 |  | 44.8 |  | 120 |  |  | LF |  | 1" 600# RF | - | 600# RF | 55 | 0 | 180 | 0 | 67.2 | 91 | 0.64
| 10 | PG-1142 | DIAPHRAGM 600# | GB-120 SUCTION | D-1-IPA-1002 | PRG | V | 0 ~ 50 | 0 | 50 | 0 ~ 50 |  | 31 |  | 94 |  |  | LF |  | 1" 600# RF | BB59C1 | 600# RF | 55 | 0 | 150 | 0 | 46.5 | 65 | 0.62
| 11 | PG-1145A | DIAPHRAGM 600# | GA-145A DISCH. | D-1-IPA-1002 | PWW | L | 0 ~ 70 | 0 | 70 | 0 ~ 70 |  | 39.25 |  | 95 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 48.1 | 0 | 150 | 0 | 58.875 | 91 | 0.56071
| 12 | PG-1145B | DIAPHRAGM 600# | GA-145B DISCH. | D-1-IPA-1002 | PWW | L | 0 ~ 70 | 0 | 70 | 0 ~ 70 |  | 39.25 |  | 95 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 48.1 | 0 | 150 | 0 | 58.875 | 91 | 0.56071
| 13 | PG-1146 | DIAPHRAGM 600# | FC-120 EQUIP. | D-1-IPA-1002 | PRG | V | 0 ~ 50 | 0 | 50 | 0 ~ 50 |  | 31 |  | 94 |  | D | LF |  | 1" 150# RF | - | HOLD | 55 | 0 | 150 | 0 | 46.5 | 65 | 0.62
| 14 | PG-1225A | DIAPHRAGM 600# | GA-225A DISCH. | D-1-IPA-2002 | PPR | L | 0 ~ 40 | 0 | 40 | 0 ~ 40 |  | 22.15 |  | 45 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 28.3 | 0 | 120 | -46 | 33.225 | 52 | 0.55375
| 15 | PG-1225B | DIAPHRAGM 600# | GA-225B DISCH. | D-1-IPA-2002 | PPR | L | 0 ~ 40 | 0 | 40 | 0 ~ 40 |  | 22.15 |  | 45 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 28.3 | 0 | 120 | -46 | 33.225 | 52 | 0.55375
| 16 | PG-1226A | DIAPHRAGM 600# | GA-226A DISCH. | D-1-IPA-2002 | PPR | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 50.9 |  | 47.2 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 58.9 | 0 | 120 | -46 | 76.35 | 104 | 0.63625
| 17 | PG-1226B | DIAPHRAGM 600# | GA-226B DISCH. | D-1-IPA-2002 | PPR | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 50.9 |  | 47.2 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 58.9 | 0 | 120 | -46 | 76.35 | 104 | 0.63625
| 18 | PG-1316A | DIAPHRAGM 600# | GA-316A DISCH. | D-1-IPA-3001 | PE | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 50.83 |  | 45 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 63.5 | 0 | 120 | 0 | 76.245 | 104 | 0.63538
| 19 | PG-1316B | DIAPHRAGM 600# | GA-316B DISCH. | D-1-IPA-3001 | PE | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 50.83 |  | 45 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 63.5 | 0 | 120 | 0 | 76.245 | 104 | 0.63538
| 20 | PG-1416A | DIAPHRAGM 600# | GA-415A DISCH. | D-1-IPA-4001 | PPF | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 49.7 | 35 | 25 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 55 | 0 | 65 | -15 | 74.55 | 104 | 0.62125
| 21 | PG-1416B | DIAPHRAGM 600# | GA-415B DISCH. | D-1-IPA-4001 | PPF | L | 0 ~ 80 | 0 | 80 | 0 ~ 80 |  | 49.7 | 35 | 25 |  | D | LF |  | 1" 600# RF | BB59C1 | 600# RF | 55 | 0 | 65 | -15 | 74.55 | 104 | 0.62125
| 22 | PG-1802 | DIAPHRAGM 600# | BFW FROM BL | D-1-IPA-8011 | BFW | L | 0 ~ 100 | 0 | 100 | 0 ~ 100 |  | 64.6 |  | 80 |  | D | LF |  | 1" 150# RF | HOLD | HOLD | 66 | 0 | 300 | 0 | 96.9 | 130 | 0.646
| 23 | PG-1845 | DIAPHRAGM 600# | FA-845A EQUIP. | D-1-IPA-8014 | NH | V | 0 ~ 110 | 0 | 110 | 0 ~ 110 |  | 70 |  | AMB |  |  | LF |  | 1" 600# RF | - | 600# RF | 77 | 0 | 65 | 0 | 105 | 143 | 0.63636
| 1 | PG-1211 | NPT | HP TO EA-211 | D-1-IPA-2001 | SH | S | 0 ~ 50 | 0 | 50 | 0 ~ 50 |  | 28 |  | 231 | S |  | LF |  | 1/2" NPT | BA59C | NPT | 46 | 0 | 400 | 0 | 42 | 65 | 0.56
| 2 | PG-1801 | NPT | HP FROM BL | D-1-IPA-8011 | HS | S | 0 ~ 70 | 0 | 70 | 0 ~ 70 |  | 40 |  | 370 | S |  | LF |  | 1/2" NPT | BA59C | NPT | 45 | F.V | 425 | 0 | 60 | 91 | 0.57143
| 3 | PG-1805 | NPT | MSS TO BG-803 | D-1-IPA-8011 | MSS | S | 0 ~ 30 | 0 | 30 | 0 ~ 30 |  | 14.6 |  | 226 | S |  | LF |  | 1/2" NPT | BB57C | NPT | 18 | 0 | 350 | 0 | 21.9 | 39 | 0.48667
| 4 | PG-1806 | NPT | MSS TO BG-803 | D-1-IPA-8011 | MSS | S | 0 ~ 30 | 0 | 30 | 0 ~ 30 |  | 14.6 |  | 226 | S |  | LF |  | 1/2" NPT | BB57C | NPT | 18 | 0 | 350 | 0 | 21.9 | 39 | 0.48667
| 5 | PG-1811 | NPT | HP TO BG-801 | D-1-IPA-8011 | HS | S | 0 ~ 70 | 0 | 70 | 0 ~ 70 |  | 40 |  | 370 | S |  | LF |  | 1/2" NPT | BA59C | NPT | 45 | F.V | 425 | 0 | 60 | 91 | 0.57143
| 6 | PG-1812 | NPT | HP TO BG-802 | D-1-IPA-8011 | HS | S | 0 ~ 70 | 0 | 70 | 0 ~ 70 |  | 40 |  | 370 | S |  | LF |  | 1/2" NPT | BA59C | NPT | 45 | F.V | 425 | 0 | 60 | 91 | 0.57143
Note on AC5: Range Max >= Design temp.
Note on AD5: 0.33<=nor/range<=0.66
Note on N29: 공정조건 변경 : 1.02->0.01