## Cover
[5 empty] | DATA SHEET | [5 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  | : M0556 |  | ANALYZER WITH HOUSE | [5 empty] | DS-IN-F-021
| PROJECT NAME |  | : IPA 50KTA Exp. | [7 empty] | REV. NO.
| LOCATION |  | : Yeosu, Korea | [7 empty] | A
| CLIENT |  | LG Chem Ltd.
| PROJECT |  | IPA 50KTA Expansion
| LOCATION |  | Yeosu, Korea
UNIT
DATA SHEET
FOR
ANALYZER WITH HOUSE
SERVEONE |  |  |  | SERVEONE CONFIDENTIAL INFORMATION: The drawings, designs, ideas, specifications and information contained in this document constitute
confidential information belonging to SERVEONE. Neither this document nor its contents may be used for any purpose or disclosed in any manner
CONSTRUCTION BUSINESS UNIT |  |  |  | without the prior written consent of SERVEONE. Reproduction of this document without the prior written consent of SERVEONE or LG Chem.,Ltd.
is prohibited. Any authorized reproduction of this document, or any part thereof, including vendors' shop drawings, shall incorporate this notice."
A |  | JUL. 02. '18 |  | FOR QUOTATION |  | J.H Jung |  | C.M Lee | K.Y Noh |  |  | H.D Yoon
Rev. |  | Date |  | Description |  | Prepared |  | Checked | Approved |  |  | Authorized
JOB NO. |  |  |  | DEPARTMENT |  |  | DOCUMENT NO. |  |  | ISSUE DATE
M0556 |  |  |  | E&I (INSTRUMENT) |  |  | DS-IN-F-021 |  |  | 2018.07.02

## Gaschro
[6 empty] | DATA SHEET | [5 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  | : M0556 |  |  | ANAYLZER (GC) | [5 empty] | DS-IN-F-021
| PROJECT NAME |  | : IPA 50KTA Exp. | [8 empty] | REV. NO.
| LOCATION |  | : Yeosu, Korea | [8 empty] | A
1 | GENERAL | Tag No. |  |  |  | AT-XXXX
2 |  | P&ID No. |  |  |  | D-1-IPA-3004
3 |  | Service |  |  |  | EA-353A/B DISCHARGE TO FA-373A/B/C
4 |  | Line Size |  | Line Class |  | 2" |  | BB55C1
5 |  | Function |  |  |  | Indicating Transmitter
6 |  | Application |  |  |  | Component Analysis
7 |  | Enclosure Class |  |  |  | Ex d IIB T4, IP56
8
9 | SERVICE CONDITION | Fluid |  | State |  | IPA (Iso Propyl Alcohol) |  |  | L
10 | [5 empty] | MAX | NOR |  | MIN | MAX | NOR |  | MIN
11 |  | Pressure |  | (kg/cm2g) |  |  | 3.5
12 |  | Temperature |  | (deg C) |  |  | 85
13 |  | S.G. @ Op |  |  |  |  | -
14 |  | Viscosity @ Op. |  | (cP) |  |  | -
15 | [5 empty] | Component | NOR (ppmv) |  | Range (ppmv) | Component | NOR (ppmv) |  | Range (ppmv)
16 |  | Component At Sample Point |  |  |  | ETOH | - |  | 0~70
17 | [5 empty] | NPA (n-PrOH) | - |  | 0~70
18 | [5 empty] | HEX (HeOH) | - |  | 0~10
19 | [5 empty] | ACETONE | - |  | 0~20
20 | [5 empty] | MeOH | - |  | 0~20
21 | [5 empty] | MIBC | - |  | 0~10
22 | [5 empty] | BENZENE | - |  | 0~10
23 | [5 empty] | H2O | - |  | 0~1000
24
25
26 | UTILITY CONDITION |  |  |  |  | MAX |  | NOR
27 |  | Steam Pressure |  |  | (kg/cm2g) | 6 |  | 3.5
28 |  | Steam Temperature |  |  | (deg C) | 200 |  | 147
29 |  | Instrument Air Pressure |  |  | (kg/cm2g) | 9 |  | 7
30 |  | Cooling Water Pressure |  |  | (kg/cm2g) | 7 |  | 5
31 |  | Cooling Water Temperature |  |  | (deg C) | 50 |  | 32
32 |  | Nitrogen Pressure |  |  | (kg/cm2g) | - |  | 7
33 | SENSOR | Type |  |  |  | TCD or FID
34 |  | Connection |  |  |  | MFR. STD
35 |  | Material |  |  |  | MFR. STD
36 |  | Process Connection |  |  |  | 1 1/2" ANSI 150# RF
37 |  | Distance From Sampling Point |  |  |  | LATER
38 |  | Electric Connection |  |  |  | 1/2" or 3/4" NPT
39 | CONTROL UNIT | Mounting Type |  |  |  | SELF STANDING
40 |  | Channel Quantity |  |  |  | MFR. STD
41 |  | Power Supply |  |  |  | AC 110V, 60Hz
42 |  | Display |  |  |  | LCD
43 |  | Alarm |  |  |  | COMMON FAULT & PURGE GAS ALARM
44 |  | Output Signal |  |  |  | 4~20mA DC
45 |  | Cycle Time |  |  |  | ≤ 5 min
46 |  | Accuracy |  |  |  | ± 1 % of F.S
47 |  | Repeatability |  |  |  | ± 1 % of F.S
48 |  | Sensibility |  |  |  | MFR. STD
49 | ACCESSORIES | Sample Probe |  |  |  | YES (316SS)
50 |  | Sample Handling System |  |  |  | YES
51 |  | Sample Return |  |  |  | TO FLARE
52 |  | Carrier Gas |  |  |  | BY VENDOR
53 |  | Calibration Gas |  |  |  | BY VENDOR
54 | PROQ. | Manufacturer
55 |  | Model
56
NOTE
| 1. | Simultaneous Detecting Type (Non-Switching) shall be supplied.
GENERAL REQUIREMENTS
| 1. | Carrier Gas : Vendor shall advise the contractor of the required type, grade, and maximum allowable contaminants
in the carrier gas for this application. (Helium is not a recommended carrier gas due to limited availability.)
A suitable cylinder storage rack, sheltered from the rain and sun shall be provided by Vendor.
All regulators, gauges, cylinder gas piping manifolds and driers, if required, shall be provided by Vendor.
| 2. | Gas Chromatographs shall be fitted with Electronic Pressure Control.
| 3. | For an individual component analysis, Automatic Peak Detection is required.
| 4. | Vendor to advise requirements of Sample Handling System.
| 5. | Process Data & Range of Individual Component will be finalized later.

## Moisture
[6 empty] | DATA SHEET | [5 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  | : M0556 |  |  | ANAYLZER (MOISTURE) | [5 empty] | DS-IN-F-021
| PROJECT NAME |  | : IPA 50KTA Exp. | [8 empty] | REV. NO.
| LOCATION |  | : Yeosu, Korea | [8 empty] | A
1 | GENERAL | Tag No. |  |  |  | AT-XXXX
2 |  | P&ID No. |  |  |  | D-1-IPA-3004
3 |  | Service |  |  |  | EA-353A/B DISCHARGE TO FA-373A/B/C
4 |  | Line Size |  | Line Class |  | 2" |  | BB55C1
5 |  | Function |  |  |  | Indicating Transmitter
6 |  | Application |  |  |  | Moisture Analysis
7 |  | Enclosure Class |  |  |  | Ex d IIB T4, IP56
8
9 | SERVICE CONDITION | Fluid |  | State |  | IPA (Iso Propyl Alcohol) |  |  | L
10 | [5 empty] | MAX | NOR |  | MIN | MAX | NOR |  | MIN
11 |  | Pressure |  | (kg/cm2g) |  |  | 3.5
12 |  | Temperature |  | (deg C) |  |  | 85
13 |  | S.G. @ Op |  |  |  |  | -
14 |  | Viscosity @ Op. |  | (cP) |  |  | -
15 | [5 empty] | Component | NOR (ppmv) |  | Range (ppmv) | Component | NOR (ppmv) |  | Range (ppmv)
16 |  | Component At Sample Point |  |  |  | H2O | - |  | 0~100
17
18
19
20
21
22
23
24
25
26 | UTILITY CONDITION |  |  |  |  | MAX |  | NOR
27 |  | Steam Pressure |  |  | (kg/cm2g) | 6 |  | 3.5
28 |  | Steam Temperature |  |  | (deg C) | 200 |  | 147
29 |  | Instrument Air Pressure |  |  | (kg/cm2g) | 9 |  | 7
30 |  | Cooling Water Pressure |  |  | (kg/cm2g) | 7 |  | 5
31 |  | Cooling Water Temperature |  |  | (deg C) | 50 |  | 32
32 |  | Nitrogen Pressure |  |  | (kg/cm2g) | - |  | 7
33 | SENSOR | Type |  |  |  | Karl-Fischer
34 |  | Connection |  |  |  | Mfr. Std
35 |  | Material |  |  |  | Mfr. Std
36 |  | Process Connection |  |  |  | 1 1/2" ANSI 150# RF
37 |  | Distance From Sampling Point |  |  |  | Later
38 |  | Electric Connection |  |  |  | 1/2" or 3/4" NPT
39 | CONTROL UNIT | Mounting Type |  |  |  | SELF STANDING
40 |  | Channel Quantity |  |  |  | Mfr. Std
41 |  | Power Supply |  |  |  | AC 110V, 60Hz
42 |  | Display |  |  |  | Digital (LCD)
43 |  | Alarm |  |  |  | Common Fault
44 |  | Output Signal |  |  |  | 4~20mA DC
45 |  | Cycle Time |  |  |  | ≤ 5 min.
46 |  | Accuracy |  |  |  | ± 1 % F.S
47 |  | Repeatability |  |  |  | ± 1 % F.S
48 |  | Sensibility |  |  |  | Mfr. Std
49 | ACCESSORIES | Sample Probe |  |  |  | Yes (316SS)
50 |  | Sample Handling System |  |  |  | Yes
51 |  | Sample Return |  |  |  | To Process
52 |  | Carrier Gas |  |  |  | -
53 |  | Calibration Gas |  |  |  | -
54 | PROQ. | Manufacturer
55 |  | Model
56
NOTE

## House
[5 empty] | DATA SHEET | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  |  | : M0556 | ANALYZER HOUSE | DS-IN-F-021
| PROJECT NAME |  |  | : IPA 50KTA Exp. |  | REV. NO.
| LOCATION |  |  | : Yeosu, Korea |  | A
1 | GENERAL | Quantity |  |  | ONE (1)
2 |  | Enclosure Class |  |  | Ex d IIB T4 (NOTE 5)
3
4
5 | EQIPMENT | Size |  |  | 3000mm(W) x 3000mm(D) x 2700mm(H)
6 |  | Structure Material |  |  | Galvanized Steel (2.3t)
7 |  | Insulation Material |  |  | Glass Wool (Note.1)
8 |  | No. of Door |  |  | 2 Doors w/ Windows, Double Glazed Safety Glass
9 |  | Door Type |  |  | Panic Hardware Self Closure Door Lock
10 |  | Extended Canopies |  |  | 1000 mm (Extending Outwards), Removable (For Transport Purpose)
11 |  | Painting |  |  | Note.2
12 |  | Lighting |  |  | Fluorescent Lights
13
14
15 | HVAC | Cooling Capacity |  |  | Mfr. Std
16 |  | Heating Capacity |  |  | Mfr. Std
17 |  | Power Supply |  |  | Note.3
18 |  | Enclosure Class |  |  | Ex d IIB T4 (NOTE 5)
19 |  | Air Intake Duct Length |  |  | 20 m w/ Fixing And Support Bracket
20 |  | Air Intake Accessories |  |  | Rain Hood, Wire Mesh Bird Screen
21
22
23 | OTHERS | Gas Detection |  |  | Combustible (2 Sets), O2 (1 Set)
24 |  | Alarm Monitoring Box |  |  | Yes
25 |  | Horn & Beacon |  |  | 1 Set (Out Side)
26 |  | Junction Box |  |  | Note.4
27 |  | Tray |  |  | Hot Dip Galv.
28 |  | Cable Gland |  |  | 316SS
29 |  | Earthing |  |  | Continuity Links Across Joints
30 |  | Utility Header |  |  | Mfr. Std
31 |  | Lifting Eye Bolt for Crane |  |  | Yes
32
33
34 | ANALYZER | Analyzer 1 Tag No. |  |  | AT-XXXX
35 |  | Analyzer 2 Tag No. |  |  | AT-XXXX
36 |  | Analyzer 3 Tag No.
37 |  | Analyzer 4 Tag No.
38 |  | Analyzer 5 Tag No.
39
40
41 | PROQ. | Manufacturer
42 |  | Model
43
44
NOTE
| 1. | Insulation
- Wall : 75t / Ceiling : 150t
| 2. | Painting
External galvanized steel walls and canopy will be primed 1 time with korbber primer RP141 for corrosion resistance protection.
inishing coat shall be 2 times korbber Enammuel RT546, final dry film thickness within 80-100 micron.
| 3. | Power
- UPS : AC 110V, 60Hz
- Lighting : AC 220V, 60Hz
- HVAC : AC 440V, 60Hz, 3Φ
| 4. | Junction Box (Power & Signal)
- Terminal : SS
- Entry : Bottom
- Spare : 20%
- According to IEC 60079-14, Stopping boxes shall be provided in the enclosure, on the wall or not more than 50 mm from
the wall of flameproof enclosures to limit the pressure piling effect and to prevent hot gases from entering the conduit system
from an enclosure containing a source of ignition.
| 5. | If hydrogen gas is used as carrier gas, enclosure class shall be Ex d IIB + Hydrogen T4.

## pH Meter
[5 empty] | DATA SHEET | [8 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  | : M0556 |  | ANALYZER (pH METER) | [8 empty] | DS-IN-F-021
| PROJECT NAME |  | : IPA 50KTA Exp. | [10 empty] | REV. NO.
| LOCATION |  | : Yeosu, Korea | [10 empty] | A
1 | GENERAL | Tag No. |  |  | AT-1134 | [11 empty] | 0
2 |  | P&ID No. |  |  | D-1-IPA-0004
3 |  | Service |  |  | EA-134 DISCHARGE
4 |  | Line Class |  |  | BB55C1
5 |  | Line Size |  | Line Schedule | 1"
6 |  | Enclosure Class |  |  | Ex d llB T4, IP56
7
8
9 | SERVICE CONDITION | Fluid |  | State | PDI |  | L
10 |  |  |  |  | MAX | NOR | MIN | MAX | NOR | MIN | MAX | NOR | MIN | MAX | NOR | MIN
11 |  | Pressure |  | (kg/cm2g) | 5.5 | 3.5
12 |  | Temperature |  | (deg C) | 45 | 40
13 |  | S.G. @ Op |  |  | 0.994
14 |  | Viscosity @ Op. |  |  | 0.69
15
16 | SENSOR | Measuring Range |  |  | 0 ~14
17 |  | Element Type |  |  | ELETRODE
18 |  | Installation Type |  |  | FLOW THROUGH
19 |  | Holder Type |  |  | CHAMBER
20 |  | Insertion Depth |  |  | MFR. STD
21 |  | Process Connection Size & Rating |  |  | 1" 150# RF
22 |  | Electriccal Connection |  |  | 1/2" NPT
23 |  | Temperature Compensation |  |  | AUTOMATIC
24 |  | Special Cable |  |  | YES
25
26 | TRANSMITTER | Display |  |  | LCD
27 |  | Output Signal |  |  | 4~20mA HART
28 |  | Accuracy |  |  | ± 0.5% of FS
29 |  | Repeatability |  |  | ± 0.5% of FS
30 |  | Mounting Type |  |  | 2" PIPE
31 |  | Electriccal Connection |  |  | 1/2" NPT
32
33 | MATERIAL | Housing |  |  | ADC
34 |  | Electrodes |  |  | MFR. STD
35 |  | Ball Valve
36 |  | Junction Box
37
38 | ACCESSORIES | Ball Valve
39 |  | Tee Adaptor
40 |  | Sampling System
41
42
43 | PROQ. | Manufacturer
44 |  | Model
45
NOTE

## Conductivity
[5 empty] | DATA SHEET | [8 empty] | CLIENT
LG Chem
DOC. NO.
| JOB NO. |  | : M0556 |  | ANALYZER (CONDUCTIVITY) | [8 empty] | DS-IN-F-021
| PROJECT NAME |  | : IPA 50KTA Exp. | [10 empty] | REV. NO.
| LOCATION |  | : Yeosu, Korea | [10 empty] | A
1 | GENERAL | Tag No. |  |  | AT-1430 | [11 empty] | 0
2 |  | P&ID No. |  |  | D-1-IPA-4003
3 |  | Service |  |  | MX-430 DISCHARGE
4 |  | Line Class |  |  | BB55C1
5 |  | Line Size |  | Line Schedule | 1"
6 |  | Enclosure Class |  |  | Ex d llB T4, IP56
7
8
9 | SERVICE CONDITION | Fluid |  | State | PCS |  | L
10 |  |  |  |  | MAX | NOR | MIN | MAX | NOR | MIN | MAX | NOR | MIN | MAX | NOR | MIN
11 |  | Pressure |  | (kg/cm2g) |  | HOLD
12 |  | Temperature |  | (deg C) |  | HOLD
13 |  | S.G. @ Op |  |  | HOLD
14 |  | Viscosity @ Op. |  |  | HOLD
15
16 | SENSOR | Measuring Range |  |  | NaOH 20 ~40 wt%
17 |  | Element Type |  |  | TOROIDAL
18 |  | Installation Type |  |  | FLOW THROUGH
19 |  | Holder Type |  |  | CHAMBER
20 |  | Insertion Depth |  |  | MFR. STD
21 |  | Process Connection Size & Rating |  |  | 1" 150# RF
22 |  | Electriccal Connection |  |  | 1/2" NPT
23 |  | Temperature Compensation |  |  | AUTOMATIC
24 |  | Special Cable |  |  | YES
25
26 | TRANSMITTER | Display |  |  | LCD
27 |  | Output Signal |  |  | 4~20mA HART
28 |  | Accuracy |  |  | ± 0.5% of FS
29 |  | Repeatability |  |  | ± 0.5% of FS
30 |  | Mounting Type |  |  | 2" PIPE
31 |  | Electriccal Connection |  |  | 1/2" NPT
32
33 | MATERIAL | Housing |  |  | ADC
34 |  | Electrodes |  |  | MFR. STD
35 |  | Ball Valve
36 |  | Junction Box
37
38 | ACCESSORIES | Ball Valve
39 |  | Tee Adaptor
40 |  | Sampling System
41
42
43 | PROQ. | Manufacturer
44 |  | Model
45
NOTE